## Introduction
Antimicrobial Resistance (AMR) Laboratory Quality Scorecard
Building AMR testing and management capacity
AMR Lab Scorecard v9.2 (Sep 2021)
Introduction
The ability to reliably isolate and identify bacterial pathogens and conduct antimicrobial susceptibility testing (AST) would enable selection of appropriate treatment leading to better patient outcomes, reduced cost and reduced antibiotic pressure for generation of antimicrobial resistance (AMR). Data from such testing would enable local and national surveillance to inform treatment guidelines and allow aggregation of data and reporting to global surveillance mechanisms such as World Health Organization (WHO) GLASS.
This structured approach to building quality-assured AMR testing and management capacity seeks to enable assessment and quality improvement of microbiology laboratories to reliably isolate and identify priority bacterial pathogens and conduct AST in urine, feces, blood, genital, pulmonary and wound samples, and to ensure the effective use of laboratory data in antimicrobial stewardship practices and management of AMR / hospital-acquired Infections (HAI) and outbreaks in health facilities.
The AMR Lab Scorecard, is integrated with the WHO-AFRO SLIPTA checklist (Version 2:2015), and consists of the following modules: (1) General AMR testing module, (2) Bacterial culture, detection, identification and AST of urine samples, (3) Bacterial culture, detection, identification and AST of fecal samples, (4) Bacterial culture, detection, identification and AST of blood, (5) Bacterial culture, detection, identification and AST of genital samples, (6) Bacterial culture, detection, identification and AST of pulmonary samples and (7) Bacterial culture, detection, identification and AST of wound samples.
The AMR Lab Quality scorecard may be used as a stand-alone scorecard to assess the quality of testing for culture, detection, identification and AST from blood, urine, fecal, genital, pulmonary and wound specimens, or can be used as part of a comprehensive SLIPTA assessment as a means to ensure detailed assessment of application of SLIPTA requirements to these test methods. The AMR Lab Scorecard technical modules are available as part of this eTool and as stand-alone hard copy checklists.
The SLIPTA checklist was developed by WHO-AFRO, in collaboration with the African Society for Laboratory Medicine (ASLM), U.S. Centers for Disease Control and Prevention (CDC) and host countries, and provides a framework for improving quality of public health laboratories in developing countries to achieve the requirements of the International Organization for Standardization (ISO)15189 standard.
Depending on the test menu of the laboratory the assessor(s) may use one (or more) of the technical scorecard modules e.g. if the microbiology laboratory performs urine culture use the "Bacterial culture, detection, identification and AST from urine samples scorecard". The SLIPTA checklist is used in conjunction with the technical scorecards. Please refer to the "A structured approach to building quality-assured AMR testing and reporting capacity- User Guide" for more details.
The "AMR Testing Scorecard eTool" supplements the hard-copy technical modules & SLIPTA checklist and automated analysis and reporting of assessments.
Acknowledgements
The development of the Guide and AMR Laboratory Scorecard was led by FIND in collaboration with and support from Becton Dickinson and Company (BD). Contributions were made by André Trollip, Heidi Albert, Cecilia Ferreyra, Cassandra Kelly-Cirino and Zachary Katz (FIND), Renuka Gadde and Courtney Maus (BD), and Tjeerd Datema and Linda Oskam (DATOS). Technical input and review were provided by John Nkengasong (Africa Centers for Disease Control and Prevention), Andrew Whitelaw (National Health Laboratory Service / University of Stellenbosch, South Africa and Beatrice Puije (African Society for Laboratory Medicine).
The AMR Laboratory Scorecard draws from a number of existing tools, including Centers for Disease Control and Prevention Laboratory Assessment of Antimicrobial Resistance Testing Capacity checklist, ICMR (India Council for Medical Research) AMR Checklist, WHO-AFRO SLIPTA checklist and FIND’s Score-TB package.

## Instructions
AMR Lab Scorecard v9.2 (Sep 2021)
Instructions
The eTool is divided into the following colour-coded TABS:
Introduction & Instructions section (Grey): this and the previous tab
Summary section (blue):
1. SLIPTA Summary Report: provides a calculated summary score for the AMR and SLIPTA assessments. Scores are automatically updated when the AMR (green) & SLIPTA (purple) tabbed sheets are completed.
2. AMR Summary Report: provides a calculated summary score for the AMR assessments (only). Scores are automatically updated when the AMR (green) tabbed sheets are completed.
AMR section (green)
1. Set Audit Scope
This TAB allows the assessor to select which technical modules will be completed. Note: Selection of the one or more technical modules is essential to using the eTool.
2. General AMR module
This TAB corresponds to the "Technical Information" section located at the beginning of all the technical modules. This section is compulsory for all assessments. The section is used to collect general technical information about the microbiology laboratory, and provides the assessor the context for performing the assessment. The "Technical Information" questions are common to all the technical modules and will only be completed once if multiple modules are used.
4. Urine module
This TAB corresponds to the "Technical Information" section located in the "Bacterial culture, detection, identification and antimicrobial susceptibility testing of urine samples" module. This section is compulsory for all laboratories testing on urine samples. Please refer to the AMR Lab Quallity Scorecard- User Guide for more details on completing this section.
5. Feces module
This TAB corresponds to the "Technical Information" section located in the "Bacterial culture, detection,identification and antimicrobial susceptibility testing of fecal samples" module. This section is compulsory for all laboratories testing on feces samples. Please refer to the AMR Lab Quallity Scorecard- User Guide for more details on completing this section.
6. Blood module
This TAB corresponds to the "Technical Information" section located in the "Bacterial culture, detection, identification and antimicrobial susceptibility testing of blood samples" module. This section is compulsory for all laboratories conducting testing on blood samples. Please refer to the AMR Lab Quallity Scorecard- User Guide for more details on completing this section.
6. Genital module
This TAB corresponds to the "Technical Information" section located in the "Bacterial culture, detection, identification and antimicrobial susceptibility testing of genital samples" module. This section is compulsory for all laboratories conducting testing on genital samples. Please refer to the AMR Lab Quallity Scorecard- User Guide for more details on completing this section.
7. Pulmonary module
This TAB corresponds to the "Technical Information" section located in the "Bacterial culture, detection, identification and antimicrobial susceptibility testing of pulmonary samples" module. This section is compulsory for all laboratories conducting testing on pulmonary samples. Please refer to the AMR Lab Quallity Scorecard- User Guide for more details on completing this section.
8. Wound module
This TAB corresponds to the "Technical Information" section located in the "Bacterial culture, detection, identification and antimicrobial susceptibility testing of wound samples" module. This section is compulsory for all laboratories conducting testing on wound samples. Please refer to the AMR Lab Quallity Scorecard- User Guide for more details on completing this section.
Pevious audit information (yellow)
This section provides place to record the scores from previous (SLIPTA and AMR) assessments conducted at the laboratory.
SLIPTA section (Purple)
To perform the SLIPTA assessment all the questions in all tabs should be completed:
1. SLIPTA aboratory profile- This TAB corresponds to the "Part I: Laboratory Profile of the SLIPTA checklist.
2. SLIPTA_S1- This TAB corresponds to the "Part II: Laboratory Audits Section 1: Documents & Records of the SLIPTA checklist.
3. SLIPTA_S2- This TAB corresponds to the "Part II: Laboratory Audits Section 2: Management Reviews of the SLIPTA checklist.
4. SLIPTA_S3- This TAB corresponds to the "Part II: Laboratory Audits Section 3: Organization & Personnel of the SLIPTA checklist.
5. SLIPTA_S4- This TAB corresponds to the "Part II: Laboratory Audits Section 4: Client Management & Customer Service of the SLIPTA checklist.
6. SLIPTA_S5- This TAB corresponds to the "Part II: Laboratory Audits Section 5: Equipment of the SLIPTA checklist.
7. SLIPTA_S6- This TAB corresponds to the "Part II: Laboratory Audits Section 6: Evaluation and Audits of the SLIPTA checklist.
8. SLIPTA_S7- This TAB corresponds to the "Part II: Laboratory Audits Section 7: Purchasing & Inventory of the SLIPTA checklist.
9. SLIPTA_S8- This TAB corresponds to the "Part II: Laboratory Audits Section 8: Process Control of the SLIPTA checklist.
10. SLIPTA_S9- This TAB corresponds to the "Part II: Laboratory Audits Section 9: Information Management of the SLIPTA checklist.
11. SLIPTA_S10- This TAB corresponds to the "Part II: Laboratory Audits Section 10: Identification of Nonconformities, Corrective and Preventative Actions of the SLIPTA checklist.
12. SLIPTA_S11- This TAB corresponds to the "Part II: Laboratory Audits Section 11: Occurrence/Incident Management & Process Improvement of the SLIPTA checklist.
13. SLIPTA_S12- This TAB corresponds to the "Part II: Laboratory Audits Section 12: Facilities and Biosafety of the SLIPTA checklist.
Completing the AMR Lab Quality Scorecard eTool:
The suggest order for completing the eTool is shown below:
The red tabbed "Set Audit Scope" sheet requires the assessor to select the AMR Modules that will be completed as part of the assessment. Selecting one or all modules, unlocks these for data entry. Note, the SLIPTA assessed tabs are always unlocked for data entry and do not require further selection.
The blue tabbed sheets display summary reporting information (once tool has been completed) and green & purple tabbed sheets are for the assessor to complete.
Questions are numbered in accordance with the AMR modules or SLIPTA checklist. Any subquestions are located below the main question and show the letter designation as found in the AMR modules or SLIPTA checklist.
Only cells colored light green are editable by the assessor. These cells either allow for a value to be selected from a drop down list or for the input of text or numbers. The drop down menu options reflect the answer choices on the AMR modules or SLIPTA checklist documents.
Scoring for each question is described in the AMR Lab Quallity Scorecard- User Guide, the spreadsheet will automatically calculate the score for each section and display the results in the summary output tabs.
If questions in the AMR modules or SLIPTA checklist documents have the option to be marked as N/A, this tool automatically adjusts the total score.

## Set Audit Scope
AMR Lab Scorecard v9.2 (Sep 2021)
Set Audit Scope
Does the laboratory perform culture and AST of urine samples?
Does the laboratory perform culture and AST of feces samples?
Does the laboratory perform culture and AST of blood samples?
Does the laboratory perform culture and AST of genital samples?
Does the laboratory perform culture and AST of pulmonary samples?
Does the laboratory perform culture and AST of wound samples?
Next step(s):
- Complete the General AMR Module (click here)
Yes
No

## General AMR Module
AMR Lab Scorecard v9.2 (Sep 2021)
General Procedures
A. General Information
Name of Assessor(s)
Title & organization of Assessor
Name of laboratory being assessed
Type of laboratory | [10 empty] | Choose from drop-down menu -->
Level of laboratory | [10 empty] | Choose from drop-down menu -->
Does the microbiology laboratory meet minimum space and infrastructure requirements? | [10 empty] | Choose from drop-down menu -->
How many hospitals and/or other health care facilities does the laboratory serve?
Location of laboratory being assessed (City/Town, County / District / Sub-district and Country)
Details of contact person at laboratory
Name
Position
Qualification
Email
Phone
Is there a clinical microbiologist and / or pathologist with experience in microbiology on staff? | [10 empty] | Choose from drop-down menu -->
If yes, how many years' experience do they have?
If no, what is the highest qualified member of laboratory staff?
Is the laboratory accredited? | [10 empty] | Choose from drop-down menu -->
If yes, name of accrediting body?
If yes, what tests is the laboratory accredited for?
Date of last assessment visit (e.g. internal audit)
- Internal
- External
Date of last SLIPTA assessment and star rating
Date of last AMR assessment visit
Date of this AMR assessment visit
Did the last assessment include assessment of bacterial culture of urine samples? | [10 empty] | Choose from drop-down menu -->
Did the last assessment include assessment of bacterial culture of feces samples? | [10 empty] | Choose from drop-down menu -->
Did the last assessment include assessment of bacterial culture of blood samples? | [10 empty] | Choose from drop-down menu -->
Did the last assessment include assessment of bacterial culture of genital samples? | [10 empty] | Choose from drop-down menu -->
Did the last assessment include assessment of bacterial culture of pulmonary samples? | [10 empty] | Choose from drop-down menu -->
Did the last assessment include assessment of bacterial culture of wound samples? | [10 empty] | Choose from drop-down menu -->
B. Technical Information
A | What procedures are available for the detection and/or identification of bacterial pathogens?
[8 empty] | Y/N | Please specify
| Conventional methods | [13 empty] | <-- Example of conventional methods: Gram stain, biochemical tests.
| Automated systems | [13 empty] | <-- Example of automated systems: Vitek, Microscan, Phoenix.
| Kit-based / serological methods | [13 empty] | <-- BioMérieux’s API identification and other similiar products.
Anaerobic methods
| Molecular detection assays (commercial) | [13 empty] | <-- Molecular detection platforms for detection and AST including real-time PCR and sequencing. Equipment may include thermocyclers, electrophoresis and gel documentation systems.
| Molecular detection assays (non-commercial/in-house) | [13 empty] | <-- Molecular detection platforms for detection and AST including real-time PCR and sequencing. Equipment may include thermocyclers, electrophoresis and gel documentation systems.
MALDI-TOF Mass spectrometry (MS)
B | What methods are available for Antimicrobial Susceptibility Tests (AST) of bacterial pathogens?
| Method |  |  |  |  | Y/N | Please specify
Automated
Gradient / disk diffusion / E-test
Disk diffusion
Manual broth dilution
| Molecular detection assays (commercial) | [15 empty] | <-- Molecular detection platforms for detection and AST including real-time PCR and sequencing. Equipment may include thermocyclers, electrophoresis and gel documentation systems.
| Molecular detection assays (non-commercial / in-house) | [15 empty] | <-- Molecular detection platforms for detection and AST including real-time PCR and sequencing. Equipment may include thermocyclers, electrophoresis and gel documentation systems.
C | Does the laboratory routinely perform AST on bacterial pathogens? | [6 empty] | Answer:
If so, which methods are used for:
|  |  |  | Automated | Gradient / Disk diffusion / E-test |  | Gradient / Disk diffusion / E-test | Manual broth dilution |  | Molecular detection assays (commercial) |  | Molecular detection assays (non-commercial / in-house)
E. coli
K. pneumoniae
Salmonella sp.
Shigella sp.
S. aureus
S. pneumoniae
A. baumannii
N. gonorrhoeae
S. pyogenes
M. catarrhalis
C. diphtheriae
H. influenzae
M. pneumoniae
Enterococcus
Enterobacteriaceae
P. aeruginosa
D | Is the following equipment available, and if so, is it functional, monitored, serviced and maintained?
[5 empty] | Available |  | Functional | Monitored |  | Serviced |  | Maintained |  | <------- | Functional: is the equipment in working order? Monitored: is the functionality of equipment regularly checked (e.g. temperature / calibrated)? Serviced: is the equipment regularly serviced by a qualified service technician? Review equipment logbook. Maintained: is the equipment regularly maintained according to the manufacturer’s recommendations (e.g. cleaning)? Review SOP and equipment logbook.
Automated blood culture instrument
| Ruler or caliper with millimeter markings | [8 empty] | NA
| Bunsen burner or micro-incinerator | [8 empty] | NA
| Wire loops for streaking | [8 empty] | NA
Turbidity meter
Microscope
| Thermometers | [8 empty] | NA
| Incubator (Aerobic) | [8 empty] | NA
| Incubator (Anaerobic) | [8 empty] | NA
| Incubator (CO2) | [8 empty] | NA
| Refrigerator (2-8°C) | [8 empty] | NA
| Freezer (-20 - -80°C) | [8 empty] | NA
Balance / scale
| Autoclave | [8 empty] | NA
Biosafety Cabinet Class II
MALDI-TOF / MS
| Molecular platforms | [12 empty] | <-- Molecular detection platforms including real-time PCR and sequencing.
Other equipment (please specify)
N/A = Not Applicable
E | How does the laboratory obtain media for bacterial culture?
[13 empty] | Urine |  | Feces |  | Blood |  | Genital | Pulmonary |  | Blood
Media is prepared on-site (non-commercial)
Media is prepared off-site (non-commercial)
|  | If Blood Agar is prepared (non-commercial), what is the source of the blood? | [20 empty] | <-- Assess the QC records to determine the quality of the blood.
Ready-made media is procured from a media supplier (commercial)
F | Which AST interpretation standard (and version) does the laboratory use (check all that apply)?
| Standard | [10 empty] | Y/N | If yes, which version is used? | [7 empty] | <-- Determine whether the laboratory has constant access, either online or offline.
Clinical & Laboratory Standards Institute (CLSI) (https://www.clsi.org)
European Committee on Antimicrobial Susceptibility Testing (EUCAST) (www.eucast.org/)
Other - please specify
G | How does the laboratory report results?
| Method |  | Y/N
Electronic
Paper
H | GLASS reporting and feedback
| Does the laboratory report cumulative pathogen & AST data to GLASS? | [8 empty] | <-- Either directly or through a supervising reference or central laboratory in case of an AMR surveillance network.
If yes, how frequently per year?
Does the laboratory receive feedback from GLASS?
If yes, how frequently per year?
Section 1: Documents and records
All generic requirements apply, see SLIPTA Section 1. In addition, assessors should review the following:
Question no. |  | Question | [9 empty] | Available score |  | Answer |  | Score |  | Comments | [5 empty] | Related to SLIPTA question
G1.1 |  | If the laboratory uses automated methods (including MS & molecular methods) for organism identification and AST (e.g. MS, molecular, Vitek, Microscan, Phoenix): | [9 empty] | 3 |  | Answer all sub questions |  | 0 | [7 empty] | 1.5
|  |  | a) Does the documentation provide instructions for preparing the inoculum in the correct medium and at the correct density? | [21 empty] | 100
|  |  | b) Does the documentation provide guidance on interpreting results generated by the software? | [21 empty] | 100
|  |  | c) Does the documentation provide guidance on how to recognize unacceptable results? | [21 empty] | 100
|  |  | d) Does the documentation outline what actions to take when unusual or unexpected AST results are documented from patient samples (e.g., reconfirm organism ID, reconfirm relevant QC, repeat testing, notify supervisor)? | [21 empty] | 100
|  |  | e) Does the documentation describe the defined QC organisms, QC frequency and expected QC results for use with the instrument? | [21 empty] | 100 | 500
G1.2 |  | If the laboratory uses kit-based methods for organism identification: e.g. BioMérieux’s API identification and other similar products. | [9 empty] | 3 |  | Answer all sub questions |  | 0 | [7 empty] | 1.5
|  |  | a) Does the documentation provide instructions for preparing the inoculum in the correct medium and at the correct density? | [21 empty] | 100
|  |  | b) Does the documentation provide guidance on interpreting results? | [21 empty] | 100
|  |  | c) Does the documentation provide guidance on how to recognize unacceptable results? | [21 empty] | 100
|  |  | d) Does the documentation describe the defined QC organisms, QC frequency and expected QC results for each test? | [21 empty] | 100 | 400
G1.3 |  | If the laboratory uses conventional methods for organism identification and AST: | [9 empty] | 3 |  | Answer all sub questions |  | 0 | [7 empty] | 1.5
|  |  | a) Does the documentation provide instructions for preparing the inoculum in the correct medium and at the correct density? | [21 empty] | 100
|  |  | b) If manual MIC methods are used, does the documentation describe specific criteria for measuring and determining the MIC endpoints? | [21 empty] | 100
|  |  | c) Does the documentation describe criteria for interpretation of the endpoint or zone size? | [21 empty] | 100
|  |  | d) Does the documentation provide guidance on how to recognize unacceptable results? | [21 empty] | 100
|  |  | e) Does the documentation describe the defined QC organisms, QC frequency and expected QC results for each test? | [21 empty] | 100 | 500
G1.4 |  | Does the laboratory provide restrictive (selective or cascade) reporting of AST? | [9 empty] | 2 |  |  |  | 0 | [7 empty] | 1.5
NOTE: TECHNICAL INFORMATION RELATED TO THIS QUESTION IS PROVIDED IN THE USER GUIDE. In cascade reporting, antimicrobial agents of each class are ranked based on a spectrum of activity, popularity or potential for the over-prescribing risk of drug resistance and cost. Thus, the reported AST should include the most appropriate and least expensive drugs, provided the organism is susceptible. Higher risk agents are only released if alternative options are lacking. In selective reporting, the susceptibilities of broad-spectrum agents and those drugs at risk for over-prescription are deliberately withheld.
Sub total | [11 empty] | 11 |  |  |  | 0
Section 2: Management Reviews
All generic requirements apply, see SLIPTA Section 2. In addition, assessors should review the following:
Question no. |  | Question | [9 empty] | Available score |  | Answer |  | Score |  | Comments | [5 empty] | Related to SLIPTA question
G2.1 |  | Does the laboratory have representation on all the following? | [9 empty] | 2 |  | Answer all sub questions |  | 0 | [7 empty] | 2.2
|  |  | - Antimicrobial Stewardship Committee / Anti-Microbial Coordination Committee | [21 empty] | 100
|  |  | - Infection Control Committee | [21 empty] | 100
|  |  | - Drug and Therapeutics Committee | [21 empty] | 100
|  |  | - Hospital Surveillance / Outbreak Team | [21 empty] | 100 | 400
G2.2 |  | Does the laboratory report findings / trends and other related important information regard bacterial culture, ID and AST results to the oversight committees regularly? | [9 empty] | 2 |  |  |  | 0 | [7 empty] | 2.2
G2.3 |  | Does the laboratory report cumulative antibiogram results for culture to the above oversight committees at least annually? | [9 empty] | 2 |  |  |  | 0 | [7 empty] | 2.2
G2.4 |  | Do laboratory management reviews include review of feedback or recommendations from the: | [9 empty] | 2 |  | Answer all sub questions |  | 0 | [7 empty] | 2.2
|  |  | - Antimicrobial Stewardship Committee / Anti-Microbial Coordination Committee | [21 empty] | 100
|  |  | - Infection Control Committee | [21 empty] | 100
|  |  | - Drug and Therapeutics Committee | [21 empty] | 100
|  |  | - Hospital Surveillance / Outbreak Team | [21 empty] | 100 | 400
Sub total | [11 empty] | 8 |  |  |  | 0
Section 4: Client Management & Customer Service
All generic requirements apply, see SLIPTA Section 4. In addition, assessors should review the following:
Question no. |  | Question | [9 empty] | Available score |  | Answer |  | Score |  | Comments | [5 empty] | Related to SLIPTA question
G4.1 |  | Does the laboratory provide feedback to clinicians (directly or via oversight committees, see G2.1) regarding: - Sample quality & sample rejection rates - Identity & frequency of isolated or identified pathogens | [9 empty] | 2 |  |  |  | 0 | [7 empty] | 4.3
Sub total | [11 empty] | 2 |  |  |  | 0
Section 5: Equipment
All generic requirements apply, see SLIPTA Section 5.
Question no. |  | Question | [9 empty] | Available score |  | Answer |  | Score |  | Comments | [5 empty] | Related to SLIPTA question
G5.1 |  | Does the laboratory use verified / validated methods for isolation / detection / identification and AST of pathogens? NOTE: TECHNICAL INFORMATION RELATED TO THIS QUESTION IS PROVIDED IN THE USER GUIDE. Includes all conventional, automated, kit-based, serological, MS and molecular (commercial & non-commercial) methods. | [9 empty] | 5 |  |  |  | 0 | [7 empty] | 5.3
G5.2 |  | Is all equipment for isolation / detection / identification and AST installed and placed in a suitable environment? | [9 empty] | 2 |  |  |  | 0 | [7 empty] | 5.1
G5.3 |  | Does the laboratory maintain all equipment for isolation / detection / identification and AST of pathogens? (see question D)? | [9 empty] | 3 |  |  |  | 0 | [7 empty] | 5.11
Sub total | [11 empty] | 10 |  |  |  | 0
Section 6: Evaluation and Audits
All generic requirements apply, see SLIPTA Section 6. Note: this checklist (with or without SLIPTA) can be used for internal and external audits.
Question no. |  | Question | [9 empty] | Available score |  | Answer |  | Score |  | Comments | [5 empty] | Related to SLIPTA question
G6.1 |  | Evaluation and audits: It is recommended that internal audits be conducted at least annually. External audits are conducted less frequently-assessors should use the recommendation of local accrediting body to determine the frequency of external audits. | [9 empty] | 5 |  | Answer all sub questions |  | 0 | [7 empty] | 6.1 & 1.5
|  |  | a) Do the laboratory policies require audits to be performed? | [21 empty] | 100
|  |  | b) Does the laboratory regularly conduct internal audits? | [21 empty] | 100
|  |  | c) Are external audits regularly conducted? | [21 empty] | 100
|  |  | d) Are audit recommendations and action plans followed up within the timeframe defined by the laboratory? | [21 empty] | 100 | 400
Sub total | [11 empty] | 5 |  |  |  | 0
Section 7: Purchasing & Inventory
All generic requirements apply, see SLIPTA Section 7. In addition, assessors should review the following:
Question no. |  | Question | [9 empty] | Available score |  | Answer |  | Score |  | Comments | [5 empty] | Related to SLIPTA question
G7.1 |  | Does the laboratory provide specifications for supplies and consumables and are they followed during the procurement process? | [9 empty] | 2 |  |  |  | 0 | [7 empty] | 7.2
G7.2 |  | Are storage areas for reagents and supplies setup, maintained and monitored according to manufacturer’s requirements? Ensure all supplies / reagents have not expired. Antibiotic packages not in use should be stored in a non-defrosting freezer, unopened and in their original packaging. Once opened, the antibiotic disks must be stored in such a way that the lot number and expiration date of each disk is always traceable. The antibiotic disk cartridges and strips should be stored in a tightly sealed container with active desiccants-the desiccants should be replaced or recharged at least monthly. | [9 empty] | 2 |  |  |  | 0 | [7 empty] | 7.8
Sub total | [11 empty] | 4 |  |  |  | 0
Section 8: Process Control
All generic requirements apply, see SLIPTA Section 8. In addition, assessors should review the following:
Question no. |  | Question | [9 empty] | Available score |  | Answer |  | Score |  | Comments | [5 empty] | Related to SLIPTA question
SPECIMEN COLLECTION
G8.1 |  | Does the Laboratory Request Form: | [9 empty] | 3 |  | Answer all sub questions |  | 0 | [7 empty] | 8.2 & 8.3
|  |  | - require the date and time of sample collection to be recorded? | [21 empty] | 100
|  |  | - have space for the date of hospital admission? | [21 empty] | 100
|  |  | - have space for the presumptive diagnosis? | [21 empty] | 100 | 300
BACTERIAL DETECTION AND / OR IDENTIFICATION
G8.2 |  | For automated, kit-based, MS or conventional methods: Refer to instructions for use for commercial and non-commercial molecular identification assays. | [9 empty] | 3 |  | Answer all sub questions |  | 0 | [7 empty] | 8.8
|  |  | Is QC performed on every new lot number/shipment of automated test reagents/ID cards/cartridges/conventional media before they are placed into use? | [21 empty] | 100
|  |  | Is the lab using inoculation medium appropriate for the procedure being performed? | [21 empty] | 100 | 200
G8.3 |  | For automated methods only: | [9 empty] | 3 |  | Answer all sub questions |  | 0 | [7 empty] | 8.10, 9.9
|  |  | Is the instrument software up to date? | [20 empty] | 9.9 | 100
|  |  | Does the laboratory confirm the detection / identification result by another method? Follow-up or confirmatory testing should be performed if the software flags a questionable result and if testing is performed using a non-commercial method (or commercial method on a non-validated sample type). | [20 empty] | 8.10a | 100 | 200
G8.4 |  | For kit-based methods only: | [9 empty] | 3 |  | Answer all sub questions |  | 0 | [7 empty] | 8.10, 7.4
|  |  | Is the manufacturer’s database used for result interpretation up to date? | [20 empty] | 7.4 | 100
|  |  | When an ID result does not reach the acceptable threshold, is there evidence that appropriate action is taken, such as repeating the test by another method or performing additional biochemical tests? | [20 empty] | 8.10a | 100 | 200
G8.5 |  | For conventional methods only: | [9 empty] | 2 |  |  |  | 0 | [7 empty] | 8.10a
When an identification result does not reach the acceptable threshold, is there evidence that appropriate action is taken, such as repeating the test by another method or performing additional biochemical tests?
BACTERIAL AST
G8.6 |  | For automated, kit-based, molecular or conventional methods: Refer to instructions for use for commercial and non-commercial molecular AST assays. NOTE: TECHNICAL INFORMATION RELATED TO THIS QUESTION IS PROVIDED IN THE USER GUIDE. | [9 empty] | 3 |  | Answer all sub questions |  | 0 | [7 empty] | 8.9
|  |  | When performing AST, are fresh isolates (<24 hours old) used? | [21 empty] | 100
|  |  | When performing AST, are well-isolated, pure colonies (as evidenced by Gram stain, colony morphology, etc.) used? | [21 empty] | 100
|  |  | Has the lab completed a QC conversion plan for all antibiotics in use? | [21 empty] | 100
|  |  | When preparing a bacterial inoculum for AST, is a 0.5 McFarland suspension used? | [21 empty] | 100
|  |  | After inoculation, are purity plates always made from the remaining suspension? | [21 empty] | 100
|  |  | Are control organisms tested with each batch of AST performed? | [21 empty] | 100 | 600
Sub total | [11 empty] | 17 |  |  |  | 0
Section 9: Information Management
All generic requirements apply, see SLIPTA Section 9. In addition, assessors should review the following:
Question no. |  | Question | [9 empty] | Available score |  | Answer |  | Score |  | Comments | [5 empty] | Related to SLIPTA question
G9.1 |  | Does the laboratory liaise with the clinical site when there is a suspected Hospital Acquired Infection (HAI) and / or nosocomial outbreak? E.g. Cluster of identical organisms isolated from a ward. | [9 empty] | 2 |  |  |  | 0 | [7 empty] | 9.3
Sub total | [11 empty] | 2 |  |  |  | 0
Section 10: Identification of Nonconformities, Corrective and Preventive Actions
All generic requirements apply, see SLIPTA Section 10.
Question no. |  | Question | [9 empty] | Available score |  | Answer |  | Score |  | Comments | [5 empty] | Related to SLIPTA question
G10.1 |  | Are all identified nonconforming activities identified and documented adequately? | [9 empty] | 5 |  |  |  | 0 | [7 empty] | 10.1
G10.2 |  | Is root cause analysis performed and corrective action taken for all non-conforming work? | [9 empty] | 3 |  |  |  | 0 | [7 empty] | 10.2
Sub total | [11 empty] | 8 |  |  |  | 0
Section 11: Occurrence/Incident Management & Process Improvement
All generic requirements apply, see SLIPTA Section 11. In addition, assessors should review the following:
Question no. |  | Question | [9 empty] | Available score |  | Answer |  | Score |  | Comments | [5 empty] | Related to SLIPTA question
G11.1 |  | Are all reports shared periodically with clinicians and AMR surveillance authorities (as applicable)? Assessors should review the guidance documents of the surveillance committees to determine the frequency that the laboratory should share its reports. If no recommendations exist, this should be at least quarterly. | [9 empty] | 2 |  |  |  | 0 | [7 empty] | 11.4 / 11.5
G11.2 |  | Do reports for clinicians and AMR surveillance authorities include at a minimum the number of samples, isolated or identified organisms and AST patterns? | [9 empty] | 2 |  |  |  | 0 | [7 empty] | 11.4 / 11.5
Sub total | [11 empty] | 4 |  |  |  | 0
Section 12: Facilities and Biosafety
All generic requirements apply, see SLIPTA Section 12. In addition, assessors should review the following:
Question no. |  | Question | [9 empty] | Available score |  | Answer |  | Score |  | Comments | [5 empty] | Related to SLIPTA question
G12.1 |  | Are the following PPE items used when processing samples? | [9 empty] | 2 |  | Answer all sub questions |  | 0 | [7 empty] | 12.16
|  |  | - Gloves | [21 empty] | 100
|  |  | - Laboratory coat | [21 empty] | 100 | 200
G12.2 |  | Waste management: | [9 empty] | 2 |  | Answer all sub questions |  | 0 | [7 empty] | 12.4
|  |  | Does the laboratory handle waste appropriately including disposal media and infectious material generated during testing? | [21 empty] | 100
|  |  | Are suitable disinfectants available for use when processing samples, are they freshly prepared, and is there evidence of their use? Clinical Microbiology Reviews, Jan. 1999, p. 147–179 | [21 empty] | 100 | 200
Sub total | [11 empty] | 4 |  |  |  | 0
Totals
[12 empty] | 75 |  |  |  | 0
Links to other AMR modules
- Urine module
- Feces module
- Blood module
- Genital module
- Pulmonary module
- Wound module
Yes
No
N/A
Yes
Partial
No
N/A
Yes
Partial
No
Public
Private
Academic
NGO
National
Reference
Provincial / County
District / Sub-district
Zonal
Field
Other

## Urine Module
AMR Lab Scorecard v9.2 (Sep 2021)
Bacterial culture, detection, identification and antimicrobial susceptibility testing of urine samples
U-A | How many urine cultures were performed last year?
It is highly recommended that assessors obtain the necessary permission to review the laboratory data. However, if assessors are unable to review the laboratory data this question is NOT compulsory for completion of the assessment.
[6 empty] | MSU |  |  |  | Catheter |  |  |  | Suprapubic
[6 empty] | Q1 | Q2 | Q3 | Q4 | Q1 | Q2 | Q3 | Q4 | Q1 | Q2 | Q3 | Q4 | Sub total | <------- | See: http://www.who.int/glass/en/ and other frequently isolated pathogens.
| Hospital-acquired | [17 empty] | <------- | Hospital-acquired infections are defined as bacterial infections in hospitalized patient (i.e. pathogenic bacterial isolated from a sample collected more than 48 hours after admission).
|  | K. pneumoniae | [15 empty] | 0
|  | E. coli | [15 empty] | 0
Other isolates
|  | - Gram positive cocci | [15 empty] | 0
|  | - Gram negative bacilli | [15 empty] | 0
|  | - Yeast | [15 empty] | 0
| Community-acquired | [17 empty] | <------- | Community-acquired infections are defined as ambulatory patients and hospitalized patients from which a sam-ple was collected less than 48 hours after admission.
|  | K. pneumoniae | [15 empty] | 0
|  | E. coli | [15 empty] | 0
Other isolates
|  | - Gram positive cocci | [15 empty] | 0
|  | - Gram negative bacilli | [15 empty] | 0
|  | - Yeast | [15 empty] | 0
| Unknown/ referred | [17 empty] | <------- | If the laboratory can’t distinguish between hospital & community acquired infections, the number of organisms isolated should be recorded as "Unknown/referred".
|  | K. pneumoniae | [15 empty] | 0
|  | E. coli | [15 empty] | 0
|  | Other isolates | [17 empty] | Total number of isolates
|  | - Gram positive cocci | [15 empty] | 0 |  | K. pneumoniae |  | 0
|  | - Gram negative bacilli | [15 empty] | 0 |  | E. coli |  | 0
|  | - Yeast | [15 empty] | 0 |  | Total other isolates: |  | 0
| TOTAL ISOLATES |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  |  | of which:
| TOTAL NUMBER OF URINE CULTURES PERFORMED | [18 empty] | - Gram pos cocci |  | 0
| TOTAL NUMBER OF CONTAMINATED URINE CULTURES | [18 empty] | - Gram neg bacilli |  | 0
| TOTAL NUMBER OF NEGATIVE URINE CULTURES | [18 empty] | - Yeast |  | 0
MSU = Mid-stream urine. Q = Quarter
U-B | How many molecular tests for urine pathogens were performed last year (Please specify)?
[5 empty] | MSU | MSU Q4: Catheter; Catheter Q4: Suprapubic; Suprapubic Q4: Sub total | <------- | Molecular tests performed on urine for the detection of urine pathogens.
[5 empty] | Q1 | Q2 | Q3 | Q4 | Q1 | Q2 | Q3 | Q4 | Q1 | Q2 | Q3 | Q4
Hospital-acquired
|  | K. pneumoniae | [14 empty] | 0
|  | E. coli | [14 empty] | 0
|  | Other isolates | [14 empty] | 0
Community-acquired
|  | K. pneumoniae | [14 empty] | 0 |  | Total number of isolates
|  | E. coli | [14 empty] | 0
|  | Other isolates | [14 empty] | 0 |  | K. pneumoniae |  | 0
| Unknown/ referred | [17 empty] | E. coli |  | 0
|  | K. pneumoniae | [14 empty] | 0 |  | Other isolates |  | 0
|  | E. coli | [14 empty] | 0
|  | Other isolates | [14 empty] | 0
| TOTAL |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
U-C | Are there any significant variations (> 20%) in the number of urine culture or other performed or organisms isolated each quarter? If ‘Yes’, please explain
[18 empty] | <------- | It is highly recommended that assessors obtain the necessary permission to review the laboratory data. However, if assessors are unable to review the laboratory this question is NOT compulsory for completion of the assessment.
Section 1: Documents and records
All generic requirements apply, see SLIPTA Section 1. In addition, assessors should review the following:
Question no. |  | Question |  |  | Suprapubic Q2: Available score; Suprapubic Q4: Answer |  | Score |  | Comments |  |  | Related to SLIPTA question
U1.1 |  | Does the laboratory have documentation covering the following processes? |  |  | Suprapubic Q2: 2; Suprapubic Q4: Answer all sub questions |  | 0 |  |  |  |  | 1.5
|  |  | a) Production of Blood Agar, MacConkey Agar or other media for urine pathogen isolation? | [20 empty] | 100
|  |  | b) Microscopic examination and urine cell count | [20 empty] | 100
|  |  | c) Processing of urine culture and molecular tests | [20 empty] | 100
|  |  | d) Detection, identification and AST of urinary pathogens | [20 empty] | 100
|  |  | e) Reporting of urine culture and molecular test results | [20 empty] | 100
|  |  | f) Interlaboratory comparison or proficiency testing (PT) | [20 empty] | 100
|  |  | g) Laboratory Safety | [20 empty] | 100 | 700
U1.2 |  | Are the documents complete, in-date and witnessed by all staff performing urine culture or molecular tests? |  |  | Suprapubic Q2: 2 |  | 0 |  |  |  |  | 1.5
See ISO15189:2012 Clause 5.5.3 for minimum requirements for a technical Standard Operating Procedure (SOP).
U1.3 |  | Are the following processes documented? |  |  | Suprapubic Q2: 3; Suprapubic Q4: Answer all sub questions |  | 0 |  |  |  |  | 1.5
|  |  | a) Rejection criteria for urine? | [20 empty] | 100
|  |  | b) Semi-quantitative culture for urine samples? | [20 empty] | 100
|  |  | c) Which organisms to identify based on relative quantities (> 10^4 CFU / ml)? | [20 empty] | 100
|  |  | d) Instructions for handling samples received after hours? | [20 empty] | 100
|  |  | e) Instructions for referral of urine culture and molecular tests for AST not performed at the laboratory? | [20 empty] | 100
|  |  | f) Turnaround time for urine culture or molecular tests? From sample collection to reporting | [20 empty] | 100
|  |  | g) Instructions on how to perform AST conversions for automated, disk diffusion, Etest / Gradient and microdilution AST? | [20 empty] | 100
|  |  | h) Definition of rare / unexpected AST results? | [20 empty] | 100
|  |  | i) Confirmatory tests for unusual or unexpected patient AST results? | [20 empty] | 100 | 900
Sub total |  |  |  |  | Suprapubic Q2: 7 |  | 0
Section 2: Management Reviews
All generic requirements apply, see SLIPTA Section 2.
Section 3: Organization & Personnel
All generic requirements apply, see SLIPTA Section 3. In addition, assessors should review the following:
Question no. |  | Question |  |  | Suprapubic Q2: Available score; Suprapubic Q4: Answer |  | Score |  | Comments |  |  | Related to SLIPTA question
U3.1 |  | Is there evidence that laboratory staff have been trained and are competent in the following: |  |  | Suprapubic Q2: 3; Suprapubic Q4: Answer all sub questions |  | 0 |  |  |  |  | 3.6
|  |  | a) Microscopic examination and urine cell count | [20 empty] | 100
|  |  | b) Processing of urine samples for culture and molecular tests | [20 empty] | 100
|  |  | c) Detection / identification and AST of urinary pathogens | [20 empty] | 100
|  |  | d) Interpretation of urine culture and molecular test results | [20 empty] | 100
|  |  | e) Reporting of urine culture and molecular test results | [20 empty] | 100
|  |  | f) QC for urine culture and molecular tests | [20 empty] | 100
|  |  | g) Laboratory safety | [20 empty] | 100 | 700
Review training records, competency assessment forms and duty rosters. Pay attention to date of training and scope of training compared with techniques being performed.
U3.2 |  | Is there evidence that laboratory staff are following the procedures described in the laboratory documentation: |  |  | Suprapubic Q2: 3; Suprapubic Q4: Answer all sub questions |  | 0 |  |  |  |  | 3.7
|  |  | a) Microscopic examination and urine cell count | [20 empty] | 100
|  |  | b) Processing of urine samples for culture and molecular tests | [20 empty] | 100
|  |  | c) Interpretation of urine culture and molecular test results | [20 empty] | 100
|  |  | d) Identification and AST of urinal pathogens | [20 empty] | 100
|  |  | e) Reporting of urine culture and molecular test results | [20 empty] | 100 | 500
Directly observe procedures being performed compared to the SOP.
Sub total |  |  |  |  | Suprapubic Q2: 6 |  | 0
Section 4: Client Management & Customer Service
All generic requirements apply, see SLIPTA Section 4. In addition, assessors should review the following:
Question no. |  | Question |  |  | Suprapubic Q2: Available score; Suprapubic Q4: Answer |  | Score |  | Comments |  |  | Related to SLIPTA question
U4.1 |  | Is there evidence that the laboratory has provided clients information / instructions on urine collection, storage and transportation to the laboratory? |  |  | Suprapubic Q2: 3 |  | 0 |  |  |  |  | 4.1
U4.2 |  | Is there evidence that the laboratory has provided clients information / instructions on interpretation of urine culture results and AST? |  |  | Suprapubic Q2: 2 |  | 0 |  |  |  |  | 4.1
Sub total |  |  |  |  | Suprapubic Q2: 5 |  | 0
Section 5: Equipment
All generic requirements apply, see SLIPTA Section 5.
Section 6: Evaluation and Audits
All generic requirements apply, see SLIPTA Section 6. Note: this checklist (with or without SLIPTA) can be used for internal and external audits.
Section 7: Purchasing & Inventory
All generic requirements apply, see SLIPTA Section 7. In addition, assessors should review the following:
Question no. |  | Question |  |  | Suprapubic Q2: Available score; Suprapubic Q4: Answer |  | Score |  | Comments |  |  | Related to SLIPTA question
U7.1 |  | Is all media for bacterial culture isolation, identification and AST stored correctly and in date (from date of manufacture media must be stored at 2-8 °C)? According to manufacturer’s requirements. |  |  | Suprapubic Q2: 2; Suprapubic Q4: Answer all sub questions |  | 0 |  |  |  |  | 7.10
|  |  | - Blood Agar | [20 empty] | 100
|  |  | - MacConkey agar | [20 empty] | 100
|  |  | - UriSelect / CLED / or equivalent | [20 empty] | 100
|  |  | - Mueller Hinton | [20 empty] | 100 | 400
Sub total |  |  |  |  | Suprapubic Q2: 2 |  | 0
Section 8: Process Control
All generic requirements apply, see SLIPTA Section 8. In addition, assessors should review the following:
Question no. |  | Question |  |  | Suprapubic Q2: Available score; Suprapubic Q4: Answer |  | Score |  | Comments |  |  | Related to SLIPTA question
SPECIMEN COLLECTION
U8.1 |  | Does the Laboratory Request Form have space to distinguish between MSU, suprapubic and urine collected from catheters? |  |  | Suprapubic Q2: 2 |  | 0 |  |  |  |  | 8.2 & 8.3
Sample types must be distinguished to allow for correct processing & disaggregation of data
U8.2 |  | If urine samples will reach the laboratory more than 2 hours post collection, are they transported to the laboratory on ice? |  |  | Suprapubic Q2: 2 |  | 0 |  |  |  |  | 8.5
MEDIA QUALITY CONTROL
U8.3 |  | Does the laboratory perform QC testing on all media before use? |  |  | Suprapubic Q2: 3; Suprapubic Q4: Answer all sub questions |  | 0 |  |  |  |  | 8.8
This includes in-house made or purchased from commercial sources
Blood agar
|  |  | Do QC records for blood agar plates demonstrate that they are checked for their ability to support growth of fastidious organisms such as S. pneumoniae? | [20 empty] | 100
|  |  | Do QC records for blood agar plates demonstrate that they are checked for their ability to show beta, alpha, and gamma hemolysis? | [20 empty] | 100
MacConkey agar (MAC)
|  |  | Do QC records for MAC plates demonstrate that they are checked for their ability to suppress growth of Gram-positive organisms while allowing the growth of Gram-negative organisms? | [20 empty] | 100
|  |  | Do QC records for MAC plates demonstrate that they are checked for their ability to allow visualization of lactose fermentation? | [20 empty] | 100
UriSelect or equivalent
|  |  | Do QC records for UriSelect or equivalent agar plates demonstrate their ability to visualize lactose fermentation and Gram-positive organisms? | [20 empty] | 100
CLED or equivalent
|  |  | Do QC records for CLED or equivalent agar plates demonstrate their ability to visualize lactose fermentation and Gram-positive organisms? | [20 empty] | 100
Mueller Hinton Agar (MHA)
|  |  | Do QC records demonstrate that MHA plates are checked for their ability to grow S. aureus & E. coli? | [20 empty] | 100 | 700
U8.4 |  | Does the laboratory: NOTE: TECHNICAL INFORMATION RELATED TO THIS QUESTION IS PROVIDED IN THE USER GUIDE. |  |  | Suprapubic Q2: 3; Suprapubic Q4: Answer all sub questions |  | 0 |  |  |  |  | 8.8
|  |  | a) Perform sterility and performance tests for every batch of culture media using certified reference strains as controls? | [20 empty] | 100
|  |  | b) Are reference strains sourced from an authorized supplier (e.g. ATCC) | [20 empty] | 100
|  |  | c) Are the reference strains stored, cultured and sub-cultured in accordance with the specification from the supplier? | [20 empty] | 100 | 300
U8.5 |  | Does the laboratory determine the cause of failed QC (root cause analysis), perform corrective actions and measure the effectiveness thereof? |  |  | Suprapubic Q2: 2 |  | 0 |  |  |  |  | 8.10
BACTERIAL URINE CULTURE PROCEDURE
U8.6 |  | Are all urine samples processed within 2 hours of collection, or a maximum of 4 hours after collection if transported on ice? |  |  | Suprapubic Q2: 2 |  | 0 |  |  |  |  | 8.5
U8.7 |  | Does the laboratory perform a cell count on all urine specimens prior to inoculation on culture media in order to determine the number of white blood cells in the urine? |  |  | Suprapubic Q2: 2 |  | 0 |  |  |  |  | 8.7
U8.8 |  | Does the laboratory have a procedure for rechecking cell counts to ensure consistency of microscopic observations/interpretations among all personnel performing microscopy (wet prep and concentrated)? |  |  | Suprapubic Q2: 2 |  | 0 |  |  |  |  | 8.7
U8.9 |  | Does the laboratory examine the urine for parasites? |  |  | Suprapubic Q2: 2 |  | 0 |  |  |  |  | 8.7
U8.10 |  | Are reference materials, such as permanent mounts, photomicrographs, NCCLS documents M15-A and M28-A2, or printed atlases available at the work bench to assist with identification of parasites? |  |  | Suprapubic Q2: 2 |  | 0 |  |  |  |  | 8.7
U8.11 |  | Does the laboratory perform a bacterial culture on all urine samples (or those with cell counts > 105 white blood cells /mL) as per their policy? |  |  | Suprapubic Q2: 2 |  | 0 |  |  |  |  | 8.7
U8.12 |  | Are MSU and urine collected from catheters plated using a calibrated 1µL loop? NOTE: TECHNICAL INFORMATION RELATED TO THIS QUESTION IS PROVIDED IN THE USER GUIDE. |  |  | Suprapubic Q2: 2 |  | 0 |  |  |  |  | 5.9
U8.13 |  | Are suprapubic urines plated using a calibrated 10µL loop? NOTE: TECHNICAL INFORMATION RELATED TO THIS QUESTION IS PROVIDED IN THE USER GUIDE. |  |  | Suprapubic Q2: 2 |  | 0 |  |  |  |  | 5.9
U8.14 |  | Which of the following media are used for primary culture of urine? NOTE: TECHNICAL INFORMATION RELATED TO THIS QUESTION IS PROVIDED IN THE USER GUIDE. |  |  | Suprapubic Q2: 2; Suprapubic Q4: Answer all sub questions |  | 0 |  |  |  |  | 8.7
|  |  | Blood Agar or equivalent | [20 empty] | 100
|  |  | MacConkey Agar or equivalent | [20 empty] | 100
|  |  | UriSelect / CLED or equivalent (UriSelect, cysteine-, lactose-, and electrolyte-deficient (CLED), if available, may replace Blood Agar and MAC) | [20 empty] | 100 | 300
U8.15 |  | Are media used for primary culture of urine incubated aerobically at 35-37°C for at least 18 hours? |  |  | Suprapubic Q2: 2 |  | 0 |  |  |  |  | 8.7
U8.16 |  | Does the lab use appropriate criteria for determining contamination of a urine culture specimen? (polymicrobial culture / no predominant colonies > 104 CFU) NOTE: TECHNICAL INFORMATION RELATED TO THIS QUESTION IS PROVIDED IN THE USER GUIDE. |  |  | Suprapubic Q2: 2 |  | 0 |  |  |  |  | 8.7
GRAM NEGATIVE BACILLI ID & AST USING CONVENTIONAL METHODS NOTE: TECHNICAL INFORMATION RELATED TO THE QUESTIONS IN THIS SECTION IS PROVIDED IN THE USER GUIDE.
U8.17 |  | Does the laboratory perform identification tests (ID) for at least the following urinary pathogens? |  |  | Suprapubic Q2: 2; Suprapubic Q4: Answer all sub questions |  | 0 |  |  |  |  | 8.7
|  |  | - K. pneumoniae | [20 empty] | 100
|  |  | - E. coli | [20 empty] | 100 | 200
U8.18 |  | Does the laboratory perform AST on at least the following urinary pathogens using an approved test method (see question F of the General AMR Module)? |  |  | Suprapubic Q2: 2; Suprapubic Q4: Answer all sub questions |  | 0 |  |  |  |  | 8.7
|  |  | K. pneumoniae | [20 empty] | 100
|  |  | E. coli | [20 empty] | 100 | 200
U8.19 |  | Is the following testing performed to identify Gram negative bacilli? |  |  | Suprapubic Q2: 2; Suprapubic Q4: Answer all sub questions |  | 0 |  |  |  |  | 8.7
|  |  | - Oxidase | [20 empty] | 100
|  |  | - Indole | [20 empty] | 100
|  |  | - Methyl Red | [20 empty] | 100
|  |  | - Voges Proskauer | [20 empty] | 100
|  |  | - Citrate | [20 empty] | 100
|  |  | - Triple Sugar Iron or Kligler Iron | [20 empty] | 100
|  |  | - Urease | [20 empty] | 100
|  |  | - Motility | [20 empty] | 100 | 800
U8.20 |  | Does the lab follow the latest CLSI /EUCAST guidelines for AST of Gram negative bacilli? (https://www.clsi.org / www.eucast.org) NOTE: TECHNICAL INFORMATION RELATED TO THIS QUESTION IS PROVIDED IN THE USER GUIDE. |  |  | Suprapubic Q2: 2 |  | 0 |  |  |  |  | 8.7
U8.21 |  | Does the laboratory use Combination Disk Test or another equivalent method for Extended Spectrum Beta-Lactamase (ESBL) screening? (Reference: J Clin Microbiol. 2013 Sep; 51(9): 2986–2990. DOI: 10.1128/JCM.00901-13) NOTE: TECHNICAL INFORMATION RELATED TO THIS QUESTION IS PROVIDED IN THE USER GUIDE. |  |  | Suprapubic Q2: 2 |  | 0 |  |  |  |  | 8.7
U8.22 |  | Does the laboratory use Combination Disk Test or another equivalent method for carbapenemase screening? NOTE: TECHNICAL INFORMATION RELATED TO THIS QUESTION IS PROVIDED IN THE USER GUIDE. |  |  | Suprapubic Q2: 2 |  | 0 |  |  |  |  | 8.7
INTERLABORATORY COMPARISON, PT AND EXTERNAL QUALITY ASSESSMENT (EQA) NOTE: TECHNICAL INFORMATION RELATED TO THE QUESTIONS IN THIS SECTION IS PROVIDED IN THE USER GUIDE.
U8.23 |  | Is the laboratory enrolled in an interlaboratory comparison or PT program for urine culture and other test organism identification, and AST? |  |  | Suprapubic Q2: 2 |  | 0 |  |  |  |  | 8.14
U8.24 |  | Did the laboratory pass the last 3 rounds of interlaboratory comparison or PT program testing? |  |  | Suprapubic Q2: 2 |  | 0 |  |  |  |  | 8.14
U8.25 |  | Does the laboratory receive onsite supervision visits as part of the EQA program for urine culture and molecular tests? |  |  | Suprapubic Q2: 2 |  | 0 |  |  |  |  | 8.14
Sub total |  |  |  |  | Suprapubic Q2: 52 |  | 0
Section 9: Information Management
All generic requirements apply, see SLIPTA Section 9. In addition, assessors should review the following:
Question no. |  | Question |  |  | Suprapubic Q2: Available score; Suprapubic Q4: Answer |  | Score |  | Comments |  |  | Related to SLIPTA question
U9.1 |  | Does the final report for bacterial urine cultures list the organisms for which the specimen was and was not cultured? The laboratory should inform the clinician on the report what organisms were excluded during the culture process. This may be either by choice of media or incubation conditions (e.g. anaerobic organisms). Assessors should review a number of laboratory reports to determine how results are reported. Procedures should be consistent with the laboratory’s SOPs. |  |  | Suprapubic Q2: 2 |  | 0 |  |  |  |  | 9.3
U9.2 |  | Does the laboratory report alert organisms which include at least? (Alert organisms are organisms with significant public health threat and / or organisms that are notifiable). |  |  | Suprapubic Q2: 2; Suprapubic Q4: Answer all sub questions |  | 0 |  |  |  |  | 9.3
|  |  | - Imipenim resistant K. pneumoniae | [20 empty] | 100
|  |  | - Carbapenem resistant Enterobacteriaceae | [20 empty] | 100
|  |  | - ESBL producing organisms | [20 empty] | 100 | 300
Sub total |  |  |  |  | Suprapubic Q2: 4 |  | 0
Section 10: Identification of Nonconformities, Corrective and Preventive Actions
All generic requirements apply, see SLIPTA Section 10.
Section 11: Occurrence/Incident Management & Process Improvement
All generic requirements apply, see SLIPTA Section 11. In addition, assessors should review the following:
Question no. |  | Question |  |  | Suprapubic Q2: Available score; Suprapubic Q4: Answer |  | Score |  | Comments |  |  | Related to SLIPTA question
U11.1 |  | Are the following performance indicators collected? It may not be possible for laboratories to distinguish between community and hospital acquired infection if this is not collected on the laboratory requisition form. |  |  | Suprapubic Q2: 3; Suprapubic Q4: Answer all sub questions |  | 0 |  |  |  |  | 11.4 & 11.5
|  |  | Number of urine culture and molecular tests performed (disaggregated by type): |  | Suprapubic Q4: Answer all sub questions | [7 empty] | 100
|  |  |  | Hospital-acquired Hospital-acquired infections are defined as bacterial infections in hospitalized patient (i.e. pathogenic bacterial isolated from a sample collected more than 48 hours after admission). | [19 empty] | 100
|  |  |  | Community-acquired Community-acquired infections are defined as ambulatory patients and hospitalized patients from which a sample was collected less than 48 hours after admission. | [19 empty] | 100
|  |  |  | Unknown / referred If the laboratory can’t distinguish between hospital & community acquired infections, the number of organisms isolated should be recorded as "Unknown/referred". | [19 empty] | 100 | 300
|  |  | Number and percentage of samples for bacterial urine culture and molecular tests rejected (disaggregated by reason e.g. leaked, insufficient volume) (target <1%) | [20 empty] | 100
|  |  | Number and percentage of urine culture with cell counts > 105 cells/ml | [20 empty] | 100
|  |  | Number of urine culture and molecular tests where pathogens were identified / isolated (disaggregated by type): |  | Suprapubic Q4: Answer all sub questions | [7 empty] | 100
|  |  |  | K. pneumoniae | [19 empty] | 100
|  |  |  | E. coli | [19 empty] | 100 | 200
|  |  | Number and percentage of contaminated urine culture tests (disaggregated by in-patient / out-patient / unknown/referred) | [20 empty] | 100
|  |  | Urine culture and molecular test TAT (disaggregated by in-patient/out-patient/unknown/referred - from sample collection to reporting) | [20 empty] | 100 | 600
Sub total |  |  |  |  | Suprapubic Q2: 3 |  | 0
Section 12: Facilities and Biosafety
All generic requirements apply, see SLIPTA Section 12.
Totals
[5 empty] | Suprapubic Q2: 79 |  | 0
Links to other AMR modules | bar: MSU Q1
- General AMR module
- Feces module
- Blood module
- Genital module
- Pulmonary module
- Wound module
Yes
No
N/A
Yes
Partial
No
N/A
Yes
Partial
No
Yes
No
N/A

## Feces Module
AMR Lab Scorecard v9.2 (Sep 2021)
Bacterial culture, detection, identification and antimicrobial susceptibility testing of fecal samples
F-A | How many feces culture and other tests were performed last year? | [12 empty] | <------- | It is highly recommended that assessors obtain the necessary permission to review the laboratory data. However, if assessors are unable to review the laboratory data this question is NOT compulsory for completion of the assessment.
[5 empty] | Feces culture |  |  |  | Molecular |  |  |  |  | <------- | Molecular tests performed on blood for the detection of bacterial fecal pathogens.
[5 empty] | Q1 | Q2 | Q3 | Q4 | Q1 | Q2 | Q3 | Q4 | Sub total | <------- | See: http://www.who.int/glass/en/ and other frequently isolated pathogens.
| Hospital-acquired | [12 empty] | <------- | Hospital-acquired infections are defined as bacterial infections in hospitalized patients (i.e. pathogenic bacterial isolated from a sample collected more than 48 hours after admission).
|  | Salmonella sp. | [10 empty] | 0
|  | Shigella sp. | [10 empty] | 0
|  | Other isolates | [10 empty] | 0
| Community-acquired | [12 empty] | <------- | Community-acquired infections are defined as ambulatory patients and hospitalized patients from which a sam-ple was collected less than 48 hours after admission.
|  | Salmonella sp. | [10 empty] | 0
|  | Shigella sp. | [10 empty] | 0
|  | Other isolates | [10 empty] | 0
| Unknown/ referred | [12 empty] | <------- | If the laboratory can’t distinguish between hospital & community acquired infections, the number of organisms isolated should be recorded as "Unknown/referred".
|  | Salmonella sp. | [10 empty] | 0
|  | Shigella sp. | [10 empty] | 0 |  | Total number of isolates per pathogen
|  | Other isolates | [10 empty] | 0
| TOTAL ISOLATES |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  |  | Salmonella sp. |  | 0
| TOTAL NUMBER OF FECES CULTURES PERFORMED | [13 empty] | Shigella sp. |  | 0
| TOTAL NUMBER OF NEGATIVE FECES CULTURES | [13 empty] | Other isolates |  | 0
Q = Quarter
F-B | Are there any significant variations (> 20%) in the number of feces culture or other tests performed or organisms isolated each quarter? If ‘Yes’, please explain
[14 empty] | <------- | It is highly recommended that assessors obtain the necessary permission to review the laboratory data. However, if assessors are unable to review the laboratory data this question is NOT compulsory for completion of the assessment.
Section 1: Documents and records
All generic requirements apply, see SLIPTA Section 1. In addition, assessors should review the following:
Question no. |  | Question | [11 empty] | Available score |  | Answer | Score | Comments | Related to SLIPTA question
F1.1 |  | Does the laboratory have documentation covering the following processes? | [11 empty] | 2 |  | Answer all sub questions | 0 |  | 1.5
|  |  | a) Production of Salmonella Shigella (SS) Agar, Selenite F broth or other media for faecal pathogen isolation? | [16 empty] | 100
|  |  | b) Microscopic examination of feces [raw] | [16 empty] | 100
|  |  | c) Processing of feces culture and molecular tests | [16 empty] | 100
|  |  | d) Detection, identification and AST of faecal pathogens | [16 empty] | 100
|  |  | e) Reporting of feces culture and molecular test results | [16 empty] | 100
|  |  | f) Interlaboratory comparison or proficiency testing (PT) | [16 empty] | 100
|  |  | g) Laboratory safety | [16 empty] | 100 | 700
F1.2 |  | Are the documents complete, in-date and witnessed by all staff performing feces culture and molecular tests? | [11 empty] | 2 |  |  | 0 |  | 1.5
See ISO15189:2012 Clause 5.5.3 for minimum requirements for a technical Standard Operating Procedure (SOP)
F1.3 |  | Are/is the following processes documented? | [11 empty] | 3 |  | Answer all sub questions | 0 |  | 1.5
|  |  | a) How to identify potential pathogens on all primary media? (SOP should describe colony appearance of potential pathogens nand define how to proceed when a potential pathogen is encountered) | [16 empty] | 100
|  |  | b) Instructions for referral of feces culture and molecular tests not performed at the laboratory? | [16 empty] | 100
|  |  | c) Instructions for handling samples received after hours? | [16 empty] | 100
|  |  | d) Instructions for referral of bacterial isolates for identification and AST? | [16 empty] | 100
|  |  | e) Instructions on how to perform AST conversions for automated, disk diffusion, Etest / Gradient and microdilution AST? | [16 empty] | 100
|  |  | f) Turnaround time for feces culture or molecular tests? From sample collection to reporting | [16 empty] | 100 | 600
Sub total | [13 empty] | 7 |  |  | 0
Section 2: Management Reviews
All generic requirements apply, see SLIPTA Section 2.
Section 3: Organization & Personnel
All generic requirements apply, see SLIPTA Section 3. In addition, assessors should review the following:
Question no. |  | Question | [11 empty] | Available score |  | Answer | Score | Comments | Related to SLIPTA question
F3.1 |  | Is there evidence that laboratory staff have been trained and are competent in the following: | [11 empty] | 3 |  | Answer all sub questions | 0 |  | 3.6
|  |  | a) Processing of feces culture and molecular tests | [16 empty] | 100
|  |  | b) Detection / identification and AST of faecal pathogens | [16 empty] | 100
|  |  | c) Interpretation of feces culture and molecular test results | [16 empty] | 100
|  |  | d) Reporting of feces culture and molecular test results | [16 empty] | 100
|  |  | e) QC for feces culture and molecular tests | [16 empty] | 100
|  |  | f) Laboratory safety | [16 empty] | 100 | 600
Review training records, competency assessment forms and duty rosters. Pay attention to date of training and scope of training compared with techniques being performed.
F3.2 |  | Is there evidence that laboratory staff are following the procedures described in the laboratory documentation? | [11 empty] | 3 |  | Answer all sub questions | 0 |  | 3.7
|  |  | a) Processing of feces culture and molecular tests | [16 empty] | 100
|  |  | b) Interpretation of feces culture and molecular test results | [16 empty] | 100
|  |  | c) Identification and AST of feces pathogens | [16 empty] | 100
|  |  | d) Reporting of feces culture and molecular test results | [16 empty] | 100 | 400
Directly observe procedures being performed compared to the SOP.
Sub total | [13 empty] | 6 |  |  | 0
Section 4: Client Management & Customer Service
All generic requirements apply, see SLIPTA Section 4. In addition, assessors should review the following:
Question no. |  | Question | [11 empty] | Available score |  | Answer | Score | Comments | Related to SLIPTA question
F4.1 |  | Is there evidence that the laboratory has provided clients information / instructions on feces collection, storage and transportation to the laboratory? | [11 empty] | 3 |  |  | 0 |  | 4.1
F4.2 |  | Is there evidence that the laboratory has provided clients information / instructions on interpretation of feces culture test results and AST? | [11 empty] | 2 |  |  | 0 |  | 4.1
Sub total | [13 empty] | 5 |  |  | 0
Section 5: Equipment
All generic requirements apply, see SLIPTA Section 5.
Section 6: Evaluation and Audits
All generic requirements apply, see SLIPTA Section 6. Note: this checklist (with or without SLIPTA) can be used for internal and external audits.
Section 7: Purchasing & Inventory
All generic requirements apply, see SLIPTA Section 7. In addition, assessors should review the following:
Question no. |  | Question | [11 empty] | Available score |  | Answer | Score | Comments | Related to SLIPTA question
F7.1 |  | Is all media for bacterial culture isolation, identification and AST stored correctly and in date (from date of manufacture media must be stored at 2-8 °C)? According to manufacturer’s requirements. | [11 empty] | 2 |  | Answer all sub questions | 0 |  | 7.10
|  |  | - SS Agar or eqivalent | [16 empty] | 100
|  |  | - Selenite F broth or equivalent | [16 empty] | 100
|  |  | - XLD or equivalent | [16 empty] | 100
|  |  | - Mueller Hinton | [16 empty] | 100 | 400
Sub total | [13 empty] | 2 |  |  | 0
Section 8: Process Control
All generic requirements apply, see SLIPTA Section 8. In addition, assessors should review the following:
Question no. |  | Question | [11 empty] | Available score |  | Answer | Score | Comments | Related to SLIPTA question
MEDIA QUALITY CONTROL
F8.1 |  | Does the laboratory perform QC testing on all media before use? This includes in-house made or purchased from commercial sources. | [11 empty] | 3 |  | Answer all sub questions | 0 |  | 8.8
SS / XLD agar or equivalent
|  |  | Do QC records demonstrate that they are checked for their ability to suppress growth of Enterobacteriaciae while allowing the growth of Salmonella & Shigella? | [16 empty] | 100
Selenite F broth or equivalent
|  |  | Do QC records demonstrate that Selenite F is checked for their ability to suppress growth of Enterobacteriaciae while allowing the growth of Salmonella & Shigella? | [16 empty] | 100
Mueller Hinton Agar (MHA)
|  |  | Do QC records demonstrate that MHA plates are checked for their ability to grow S. aureus & E. coli? | [16 empty] | 100 | 300
F8.2 |  | Does the laboratory: NOTE: TECHNICAL INFORMATION RELATED TO THIS QUESTION IS PROVIDED IN THE USER GUIDE. | [11 empty] | 3 |  | Answer all sub questions | 0 |  | 8.8
|  |  | a) Perform sterility and performance tests for every batch of culture media using certified reference strains as controls? | [16 empty] | 100
|  |  | b) Are reference strains sourced from an authorized supplier (e.g. ATCC)? | [16 empty] | 100
|  |  | c) Are the reference strains stored, cultured and sub-cultured in accordance with the specification from the supplier? | [16 empty] | 100 | 300
F8.3 |  | Does the laboratory determine the cause of failed QC (root cause analysis), perform corrective actions and measure the effectiveness thereof? | [11 empty] | 2 |  |  | 0 |  | 8.10
BACTERIAL FECES CULTURE PROCEDURE
F8.4 |  | Does the laboratory perform a microscopic examination [wet prep & concentrated] for parasites on all feces specimens? | [11 empty] | 2 |  |  | 0 |  | 8.5
F8.5 |  | Are reference materials, such as permanent mounts, photomicrographs, NCCLS documents M15-A and M28-A2, or printed atlases available at the work bench to assist with identification of parasites? | [11 empty] | 2 |  |  | 0 |  | 8.7
F8.6 |  | Does the laboratory procedure include rechecking microscopic observations/interpretations among all personnel performing microscopy (wet prep and concentrated)? | [11 empty] | 2 |  |  | 0 |  | 8.7
F8.7 |  | Are the following media used for primary culture of feces? Either SS agar or XLD agar or equivalent and Selenite F broth or equivalent. NOTE: TECHNICAL INFORMATION RELATED TO THIS QUESTION IS PROVIDED IN THE USER GUIDE. | [11 empty] | 2 |  | Answer all sub questions | 0 |  | 8.7
|  |  | SS Agar or equivalent | [16 empty] | 100
|  |  | Selenite F broth or equivalent | [16 empty] | 100
|  |  | XLD agar or equivalent | [16 empty] | 100 | 300
F8.8 |  | In cholera endemic areas, does the laboratory perform an isolation procedure for Vibrio using TCBS and alkaline peptone water? | [11 empty] | 2 |  |  | 0 |  | 8.7
F8.9 |  | Does the feces processing procedure include plating from Selenite F broth or equivalent to SS / XLD agar or equivalent? | [11 empty] | 2 |  |  | 0 |  | 8.7
F8.10 |  | Are media used for primary culture of feces incubated at 35-37°C for at least 18 hours? | [11 empty] | 2 |  |  | 0 |  | 8.7
F8.11 |  | Are media used for primary culture of feces incubated aerobically? | [11 empty] | 2 |  |  | 0 |  | 8.7
SALMONELLA & SHIGELLA IDENTIFICATION AND AST BY CONVENTIONAL METHODS NOTE: TECHNICAL INFORMATION RELATED TO THE QUESTIONS IN THIS SECTION IS PROVIDED IN THE USER GUIDE.
F8.12 |  | Does the laboratory perform identification tests (ID) for at least the following enteric pathogens? | [11 empty] | 2 |  | Answer all sub questions | 0 |  | 8.7
|  |  | Salmonella sp. | [16 empty] | 100
|  |  | Shigella sp. | [16 empty] | 100 | 200
F8.13 |  | Does the laboratory perform AST on at least the following enteric pathogens using an approved test method? | [11 empty] | 2 |  | Answer all sub questions | 0 |  | 8.7
|  |  | Salmonella sp. | [16 empty] | 100
|  |  | Shigella sp. | [16 empty] | 100 | 200
F8.14 |  | Is the following testing performed for Salmonella & Shigella identification: | [11 empty] | 2 |  | Answer all sub questions | 0 |  | 8.7
|  |  | Oxidase | [16 empty] | 100
|  |  | Indole | [16 empty] | 100
|  |  | Methyl Red | [16 empty] | 100
|  |  | Voges Proskauer | [16 empty] | 100
|  |  | Citrate | [16 empty] | 100
|  |  | Triple Sugar Iron or Kligler Iron | [16 empty] | 100
|  |  | Urease | [16 empty] | 100
|  |  | Motility | [16 empty] | 100
|  |  | Shigella serology | [16 empty] | 100
|  |  | Salmonella serology | [16 empty] | 100 | 1000
F8.15 |  | Is Salmonella AST done as per current CLSI/EUCAST guidelines? (https://www.clsi.org / www.eucast.org) NOTE: TECHNICAL INFORMATION RELATED TO THIS QUESTION IS PROVIDED IN THE USER GUIDE. | [11 empty] | 2 |  |  | 0 |  | 8.7
F8.16 |  | Is Shigella AST done as per current CLSI/EUCAST guidelines? (https://www.clsi.org / www.eucast.org) NOTE: TECHNICAL INFORMATION RELATED TO THIS QUESTION IS PROVIDED IN THE USER GUIDE. | [11 empty] | 2 |  |  | 0 |  | 8.7
F8.17 |  | Does the laboratory use Combination Disk Test or another equivalent method for Extended Spectrum Beta-Lactamase (ESBL) screening? (J Clin Microbiol. 2013 Sep; 51(9): 2986–2990) NOTE: TECHNICAL INFORMATION RELATED TO THIS QUESTION IS PROVIDED IN THE USER GUIDE. | [11 empty] | 2 |  |  | 0 |  | 8.7
F8.18 |  | Does the laboratory use Combination Disk Test or another equivalent method for carbapenemase screening? NOTE: TECHNICAL INFORMATION RELATED TO THIS QUESTION IS PROVIDED IN THE USER GUIDE. | [11 empty] | 2 |  |  | 0 |  | 8.7
INTERLABORATORY COMPARISON, PT AND EXTERNAL QUALITY ASSESSMENT (EQA) NOTE: TECHNICAL INFORMATION RELATED TO THE QUESTIONS IN THIS SECTION IS PROVIDED IN THE USER GUIDE.
F8.19 |  | Is the laboratory enrolled in an interlaboratory comparison or PT program for feces and molecular tests for organism identification and AST? | [11 empty] | 2 |  |  | 0 |  | 8.14
F8.20 |  | Did the laboratory pass the last 3 rounds of interlaboratory comparison or PT program testing? | [11 empty] | 2 |  |  | 0 |  | 8.14
F8.21 |  | Does the laboratory receive onsite supervision visits as part of the EQA program for feces culture and molecular tests? | [11 empty] | 2 |  |  | 0 |  | 8.14
Sub total | [13 empty] | 44 |  |  | 0
Section 9: Information Management
All generic requirements apply, see SLIPTA Section 9. In addition, assessors should review the following:
Question no. |  | Question | [11 empty] | Available score |  | Answer | Score | Comments | Related to SLIPTA question
F9.1 |  | Does the final report for feces culture list the organisms for which the specimen was and was not cultured? The laboratory should inform the clinician on the report what organisms were excluded during the culture process. This may be either by choice of media or incubation conditions (e.g. anaerobic organisms). Assessors should review a number of laboratory reports to determine how results are reported. Procedures should be consistent with the laboratory’s SOPs. | [11 empty] | 2 |  |  | 0 |  | 9.3
F9.2 |  | Does the laboratory report alert organisms which include at least: (Alert organisms are organisms with significant public health threat and/or organisms that are notifiable). | [11 empty] | 2 |  | Answer all sub questions | 0 |  | 9.3
|  |  | ESBL producing organisms | [16 empty] | 100
|  |  | Carbapenem resistant Salmonella & Shigella | [16 empty] | 100 | 200
Sub total | [13 empty] | 4 |  |  | 0
Section 10: Identification of Nonconformities, Corrective and Preventive Actions
All generic requirements apply, see SLIPTA Section 10.
Section 11: Occurrence/Incident Management & Process Improvement
All generic requirements apply, see SLIPTA Section 11. In addition, assessors should review the following:
Question no. |  | Question | [11 empty] | Available score |  | Answer | Score | Comments | Related to SLIPTA question
F11.1 |  | Are the following performance indicators collected? It may not be possible for laboratories to distinguish between community and hospital acquired infection if this is not collected on the laboratory requisition form | [11 empty] | 3 |  | Answer all sub questions | 0 |  | 11.4 & 11.5
|  |  | Number of feces culture and molecular tests performed (disaggregated by type): | [12 empty] | Answer all sub questions |  |  |  | 100
|  |  |  | Hospital-acquired (bacterial infections in hospitalized patient (i.e. pathogenic bacterial isolated from a sample collected more than 48 hours after admission)) | [15 empty] | 100
|  |  |  | Community-acquired (ambulatory patients and hospitalized patients from which a sample was collected less than 48 hours after admission) | [15 empty] | 100
|  |  |  | Unknown / referred If the laboratory can’t distinguish between hospital & community acquired infections, the number of organisms isolated should be recorded as "Unknown/referred" | [15 empty] | 100 | 300
|  |  | Number and percentage of feces culture tests rejected (disaggregated by reason e.g. leaked, insufficient volume) (target <1%) | [16 empty] | 100
|  |  | Number and percentage of feces culture tests where parasites were observed | [16 empty] | 100
|  |  | Number of feces culture tests where pathogens were isolated or identified (disaggregated by type): | [12 empty] | Answer all sub questions |  |  |  | 100
|  |  |  | Salmonella sp. | [15 empty] | 100
|  |  |  | Shigella sp. | [15 empty] | 100
|  |  |  | V. cholerae | [15 empty] | 100 | 300
|  |  | Feces culture and molecular tests TAT (disaggregated by in-patient/out-patient/unknown/referred - from sample collection to reporting) | [16 empty] | 100 | 500
Sub total | [13 empty] | 3 |  |  | 0
Section 12: Facilities and Biosafety
All generic requirements apply, see SLIPTA Section 12.
Totals
[14 empty] | 71 |  |  | 0
Links to other AMR modules | bar: Feces culture Q1
- General AMR module
- Urine module
- Blood module
- Genital module
- Pulmonary module
- Wound module
Yes
No
Yes
Partial
No
N/A
Yes
Partial
No
Yes
No
N/A

## Blood Module
AMR Lab Scorecard v9.2 (Sep 2021)
Bacterial culture, detection, identification and antimicrobial susceptibility testing of blood samples
B-A | What blood culture system does the laboratory use?
| Automated | [10 empty] | Type:
Manual
B-B | How many blood culture and other tests were performed last year? | [16 empty] | <------- | It is highly recommended that assessors obtain the necessary permission to review the laboratory data. However, if assessors are unable to review the laboratory data this question is NOT compulsory for completion of the assessment.
[5 empty] | Manual |  |  |  | Automated |  |  |  | Molecular |  |  |  |  | <------- | Molecular tests performed on blood for the detection of bacterial blood pathogens.
[5 empty] | Q1 | Q2 | Q3 | Q4 | Q1 | Q2 | Q3 | Q4 | Q1 | Q2 | Q3 | Q4 | Sub total | <------- | See: http://www.who.int/glass/en/ and other frequently isolated pathogens.
| Hospital-acquired | [16 empty] | <------- | Hospital-acquired infections are defined as bacterial infections in hospitalized patients (i.e. pathogenic bacterial isolated from a sample collected more than 48 hours after admission).
|  | S. aureus | [14 empty] | 0
|  | Coagulase-negative Staphylococcus | [14 empty] | 0
|  | S. pneumoniae | [14 empty] | 0
|  | Enterococcus sp. | [14 empty] | 0
|  | E. coli | [14 empty] | 0
|  | K. pneumoniae | [14 empty] | 0
|  | A. baumannii | [14 empty] | 0
|  | Salmonella sp. | [14 empty] | 0
Other isolates
|  | - Gram positive cocci | [14 empty] | 0
|  | - Gram negative bacilli | [14 empty] | 0
|  | - Yeast | [14 empty] | 0
| Community-acquired | [16 empty] | <------- | Community-acquired infections are defined as ambulatory patients and hospitalized patients from which a sample was collected less than 48 hours after admission.
|  | S. aureus | [14 empty] | 0
|  | Coagulase-negative Staphylococcus | [14 empty] | 0
|  | S. pneumoniae | [14 empty] | 0
|  | Enterococcus sp. | [14 empty] | 0
|  | E. coli | [14 empty] | 0
|  | K. pneumoniae | [14 empty] | 0
|  | A. baumannii | [14 empty] | 0
|  | Salmonella sp. | [14 empty] | 0
Other isolates
|  | - Gram positive cocci | [14 empty] | 0
|  | - Gram negative bacilli | [14 empty] | 0
|  | - Yeast | [14 empty] | 0
| Unknown/referred | [16 empty] | <------- | If the laboratory can’t distinguish between hospital & community acquired infections, the number of organisms isolated should be recorded as "Unknown/referred".
|  | S. aureus | [14 empty] | 0
|  | Coagulase-negative Staphylococcus | [14 empty] | 0 |  | Total number of isolates per pathogen
|  | S. pneumoniae | [14 empty] | 0
|  | Enterococcus sp. | [14 empty] | 0 |  | S. aureus |  | 0
|  | E. coli | [14 empty] | 0 |  | S. pneumoniae |  | 0
|  | K. pneumoniae | [14 empty] | 0 |  | Coagulase-negative Staphylococcus |  | 0
|  | A. baumannii | [14 empty] | 0
|  | Salmonella sp. | [14 empty] | 0 |  | E. coli |  | 0
|  | Other isolates | [16 empty] | Enterococcus sp. |  | 0
|  | - Gram positive cocci | [14 empty] | 0 |  | K. pneumoniae |  | 0
|  | - Gram negative bacilli | [14 empty] | 0 |  | A. baumannii |  | 0
|  | - Yeast | [14 empty] | 0 |  | Salmonella sp. |  | 0
| TOTAL ISOLATES |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | Total other isolates, of which: |  | 0
| TOTAL NUMBER OF BLOOD CULTURES PERFORMED | [17 empty] | - Gram positive cocci |  | 0
| TOTAL NUMBER OF CONTAMINATED BLOOD CULTURES | [17 empty] | - Gram negative bacilli |  | 0 | bar: Molecular Q1 to Molecular Q4
| TOTAL NUMBER OF NEGATIVE BLOOD CULTURES | [17 empty] | - Yeast |  | 0
Q = Quarter
B-C | Are there any significant variations (> 20%) in the number of blood culture or other tests performed or organisms isolated or identified each quarter? If ‘Yes’, please explain
[18 empty] | <------- | It is highly recommended that assessors obtain the necessary permission to review the laboratory data. However, if assessors are unable to review the laboratory data this question is NOT compulsory for completion of the assessment.
Section 1: Documents and records
All generic requirements apply, see SLIPTA Section 1. In addition, assessors should review the following:
Question no. |  | Question |  |  | Molecular Q2: Available score; Molecular Q4: Answer |  | Score |  | Comments |  | Related to SLIPTA question
B1.1 |  | Does the laboratory have documentation covering the following processes? |  |  | Molecular Q2: 2; Molecular Q4: Answer all sub questions |  | 0 |  |  |  | 1.5
|  |  | a) Production of Blood Agar, MacConkey Agar or other media for blood culture pathogen isolation? | [19 empty] | 100
|  |  | b) Processing of blood samples | [19 empty] | 100
|  |  | c) Detection, identification and AST of blood pathogens | [19 empty] | 100
|  |  | d) Reporting of blood culture and molecular test results | [19 empty] | 100
|  |  | e) Interlaboratory comparison or proficiency testing (PT) | [19 empty] | 100
|  |  | f) Laboratory safety | [19 empty] | 100 | 600
B1.2 |  | Are the documents complete, in-date and witnessed by all staff performing blood and molecular tests? |  |  | Molecular Q2: 2 |  | 0 |  |  |  | 1.5
See ISO15189:2012 Clause 5.5.3 for minimum requirements for a technical Standard Operating Procedure (SOP).
B1.3 |  | Are/is the following processes/information documented? |  |  | Molecular Q2: 3; Molecular Q4: Answer all sub questions |  | 0 |  |  |  | 1.5
|  |  | a) Rejection criteria for blood? | [19 empty] | 100
|  |  | b) A policy for reporting critical results? | [19 empty] | 100
|  |  | c) Procedure for immediate reporting of Gram stain results of positive blood cultures? | [19 empty] | 100
|  |  | d) Instructions for reporting blood culture tests with mixed bacterial growth? | [19 empty] | 100
|  |  | e) Instructions for referral of blood culture or molecular tests at the laboratory? | [19 empty] | 100
|  |  | f) Instructions for handling samples received after hours? | [19 empty] | 100
|  |  | g) Instructions for referral of bacterial isolates for identification and AST? | [19 empty] | 100
|  |  | h) Instructions on how to perform AST conversions for automated, disk diffusion, Etest / Gradient and microdilution AST? | [19 empty] | 100
|  |  | i) Turnaround time for blood culture or molecular tests? From sample collection to reporting. | [19 empty] | 100
|  |  | J) Definition of rare / unexpected AST results? | [19 empty] | 100
|  |  | k) Confirmatory tests for unusual or unexpected patient AST results? | [19 empty] | 100 | 1100
Sub total |  |  |  |  | Molecular Q2: 7 |  | 0
Section 2: Management Reviews
All generic requirements apply, see SLIPTA Section 2.
Section 3: Organization & Personnel
All generic requirements apply, see SLIPTA Section 3. In addition, assessors should review the following:
Question no. |  | Question |  |  | Molecular Q2: Available score; Molecular Q4: Answer |  | Score |  | Comments |  | Related to SLIPTA question
B3.1 |  | Is there evidence that laboratory staff have been trained and are competent in the following: |  |  | Molecular Q2: 3; Molecular Q4: Answer all sub questions |  | 0 |  |  |  | 3.6
|  |  | a) Processing of blood culture and other tests | [19 empty] | 100
|  |  | b) Identification and AST of blood pathogens | [19 empty] | 100
|  |  | c) Interpretation of bacterial blood culture and molecular tests | [19 empty] | 100
|  |  | d) Reporting of blood culture and molecular test results | [19 empty] | 100
|  |  | e) QC, EQA & PT for blood culture and molecular tests | [19 empty] | 100
|  |  | f) Laboratory safety | [19 empty] | 100 | 600
Review training records, competency assessment forms and duty rosters. Pay attention to date of training and scope of training compared with techniques being performed.
B3.2 |  | Is there evidence that laboratory staff are following the procedures described in the laboratory documentation: |  |  | Molecular Q2: 3; Molecular Q4: Answer all sub questions |  | 0 |  |  |  | 3.7
|  |  | a) Processing of blood for culture and molecular tests | [19 empty] | 100
|  |  | b) Interpretation of blood culture results | [19 empty] | 100
|  |  | c) Identification and AST of blood pathogens | [19 empty] | 100
|  |  | d) Reporting of blood culture and molecular test results | [19 empty] | 100 | 400
Directly observe procedures being performed compared to the SOP.
Sub total |  |  |  |  | Molecular Q2: 6 |  | 0
Section 4: Client Management & Customer Service
All generic requirements apply, see SLIPTA Section 4. In addition, assessors should review the following:
Question no. |  | Question |  |  | Molecular Q2: Available score; Molecular Q4: Answer |  | Score |  | Comments |  | Related to SLIPTA question
B4.1 |  | Is there evidence that the laboratory has provided clients information / instructions on blood sample collection, storage and transportation to the laboratory? Does the information / instructions include: |  |  | Molecular Q2: 3; Molecular Q4: Answer all sub questions |  | 0 |  |  |  | 4.1
|  |  | a) Use of sterile techniques for drawing and handling of blood cultures? | [19 empty] | 100
|  |  | b) Recommendations for the appropriate volume of blood per culture? In case of automated blood culture, the volume should be consistent with the manufacturer’s instruction for use. | [19 empty] | 100
|  |  | c) Collection procedures for culture of anaerobic organisms? | [19 empty] | 100
|  |  | d) Collection procedures for blood cultures on pediatric patients? | [19 empty] | 100
|  |  | e) Interpretation of contaminated results? | [19 empty] | 100
|  |  | f) Frequency of sampling for blood culture? | [19 empty] | 100 | 600
B4.2 |  | Is there evidence that the laboratory has provided clients information / instructions on interpretation of blood culture results and AST? |  |  | Molecular Q2: 2 |  | 0 |  |  |  | 4.1
Sub total |  |  |  |  | Molecular Q2: 5 |  | 0
Section 5: Equipment
All generic requirements apply, see SLIPTA Section 5.
Section 6: Evaluation and Audits
All generic requirements apply, see SLIPTA Section 6. Note: this checklist (with or without SLIPTA) can be used for internal and external audits.
Section 7: Purchasing & Inventory
All generic requirements apply, see SLIPTA Section 7. In addition, assessors should review the following:
Question no. |  | Question |  |  | Molecular Q2: Available score; Molecular Q4: Answer |  | Score |  | Comments |  | Related to SLIPTA question
B7.1 |  | Is all media for bacterial culture isolation, identification and AST stored correctly and in date (from date of manufacture media must be stored at 2-8 °C)? According to manufacturer’s requirements. |  |  | Molecular Q2: 2; Molecular Q4: Answer all sub questions |  | 0 |  |  |  | 7.10
|  |  | - Blood Agar | [19 empty] | 100
|  |  | - MacConkey agar | [19 empty] | 100
|  |  | - Mueller Hinton | [19 empty] | 100 | 300
Sub total |  |  |  |  | Molecular Q2: 2 |  | 0
Section 8: Process Control
All generic requirements apply, see SLIPTA Section 8. In addition, assessors should review the following:
Question no. |  | Question |  |  | Molecular Q2: Available score; Molecular Q4: Answer |  | Score |  | Comments |  | Related to SLIPTA question
MEDIA QUALITY CONTROL
B8.1 |  | Does the laboratory perform QC testing on all media before use? This includes in-house made or purchased from commercial sources. |  |  | Molecular Q2: 3; Molecular Q4: Answer all sub questions |  | 0 |  |  |  | 8.8
Blood agar
|  |  | Do QC records for blood agar plates demonstrate that they are checked for their ability to support growth of fastidious organisms such as S. pneumoniae? | [19 empty] | 100
|  |  | Do QC records for blood agar plates demonstrate that they are checked for their ability to show beta, alpha, and gamma hemolysis? | [19 empty] | 100
MacConkey agar (MAC)
|  |  | Do QC records for MAC plates demonstrate that they are checked for their ability to suppress growth of Gram-positive organisms while allowing the growth of Gram-negative organisms? | [19 empty] | 100
|  |  | Do QC records for MAC plates demonstrate that they are checked for their ability to allow visualization of lactose fermentation? | [19 empty] | 100
Mueller Hinton Agar (MHA)
|  |  | Do QC records demonstrate that MHA plates are checked for their ability to grow S. aureus & E. coli? | [19 empty] | 100 | 500
B8.2 |  | Does the laboratory: NOTE: TECHNICAL INFORMATION RELATED TO THIS QUESTION IS PROVIDED IN THE USER GUIDE. |  |  | Molecular Q2: 2; Molecular Q4: Answer all sub questions |  | 0 |  |  |  | 8.8
|  |  | a) Perform sterility and performance tests for every batch of culture media using certified reference strains as controls? | [19 empty] | 100
|  |  | b) Are reference strains sourced from an authorized supplier (e.g. ATCC)? | [19 empty] | 100
|  |  | c) Are the reference strains stored, cultured and sub-cultured in accordance with the specification from the supplier? | [19 empty] | 100 | 300
B8.3 |  | Does the laboratory determine the cause of failed QC (root cause analysis), perform corrective actions and measure the effectiveness thereof? |  |  | Molecular Q2: 2 |  | 0 |  |  |  | 8.10
BACTERIAL BLOOD CULTURE PROCEDURE NOTE: TECHNICAL INFORMATION RELATED TO THE QUESTIONS IN THIS SECTION IS PROVIDED IN THE USER GUIDE.
For complete recommended procedure, see the User Guide.
B8.4 |  | Is blood incubated for a minimum of 5 days before being discarded if there is no visible sign of organism growth? |  |  | Molecular Q2: 2 |  | 0 |  |  |  | 8.5
B8.5 |  | Are incubating blood cultures visually examined each day for signs of growth (e.g. turbidity, hemolysis or gas production)? |  |  | Molecular Q2: 2 |  | 0 |  |  |  | 8.7
BACTERIAL BLOOD CULTURE PROCEDURE (WORK-UP) NOTE: TECHNICAL INFORMATION RELATED TO THE QUESTIONS IN THIS SECTION IS PROVIDED IN THE USER GUIDE.
B8.6 |  | Are Gram stains performed for all blood cultures showing any sign of positive growth (e.g. turbidity, hemolysis, or gas production?) |  |  | Molecular Q2: 2 |  | 0 |  |  |  | 8.7
B8.7 |  | Is sub-culture of positive primary blood cultures done based on Gram stain result? |  |  | Molecular Q2: 2 |  | 0 |  |  |  | 8.7
B8.8 |  | Are blood culture subculture plates incubated at 35-37°C? |  |  | Molecular Q2: 2 |  | 0 |  |  |  | 8.7
B8.9 |  | Does the laboratory report blood cultures as contaminated if they contain organisms that should be considered contaminants? (e.g. Bacillus sp., Coagulase-negative Staphylococcus, Corynebacterium sp.) |  |  | Molecular Q2: 2 |  | 0 |  |  |  | 8.7
B8.10 |  | Are blood culture bottles which showed signs of positive growth, but from which no aerobic bacteria were isolated, sub-cultured to chocolate agar? |  |  | Molecular Q2: 2 |  | 0 |  |  |  | 8.7
STAPHYLOCOCCUS SP. ID & AST BY CONVENTIONAL METHODS NOTE: TECHNICAL INFORMATION RELATED TO THE QUESTIONS IN THIS SECTION IS PROVIDED IN THE USER GUIDE.
B8.11 |  | Is the following testing performed for S. aureus identification? If the laboratory performs penicillin AST, it is recommended that S. aureus isolates with penicillin zones sizes or MICs in the susceptible range are tested for B-lactamase production using the zone-edge test or a nitrocefin test before being reported as penicillin susceptible. |  |  | Molecular Q2: 2; Molecular Q4: Answer all sub questions |  | 0 |  |  |  | 8.7
|  |  | Catalase | [19 empty] | 100
|  |  | Coagulase (slide or tube) | [19 empty] | 100
|  |  | Mannitol Salt Agar (MSA) | [19 empty] | 100
|  |  | Dnase | [19 empty] | 100 | 400
B8.12 |  | Does S. aureus AST include the following antibiotics: NOTE: TECHNICAL INFORMATION RELATED TO THIS QUESTION IS PROVIDED IN THE USER GUIDE. If oxacillin and cefoxitin results are discrepant for S. aureus (one is susceptible and one is resistant), the laboratory should repeat the testing. Note: oxacillin testing should always be tested by MIC (not disc diffusion). If the results remain discrepant, oxacillin should be reported as resistant. |  |  | Molecular Q2: 2; Molecular Q4: Answer all sub questions |  | 0 |  |  |  | 8.7
|  |  | Cefoxitin | [19 empty] | 100
|  |  | Vancomycin | [19 empty] | 100 | 200
B8.13 |  | Does the laboratory detect methicillin/nafcillin resistance in S. aureus using oxacillin disk or another? NOTE: TECHNICAL INFORMATION RELATED TO THIS QUESTION IS PROVIDED IN THE USER GUIDE. |  |  | Molecular Q2: 2 |  | 0 |  |  |  | 8.7
STREPTOCOCCUS SP. ID & AST BY CONVENTIONAL METHODS
B8.14 |  | Is the following testing performed for Streptococcus sp. identification? |  |  | Molecular Q2: 2; Molecular Q4: Answer all sub questions |  | 0 |  |  |  | 8.7
|  |  | Bacitracin | [19 empty] | 100
|  |  | Pyrrolidonyl Arylamidase (PYR) | [19 empty] | 100
|  |  | Bile solubility | [19 empty] | 100
|  |  | Optochin | [19 empty] | 100
|  |  | S. pneumoniae latex | [19 empty] | 100 | 500
B8.15 |  | Does Streptococcus sp. AST include the following antibiotics: NOTE: TECHNICAL INFORMATION RELATED TO THIS QUESTION IS PROVIDED IN THE USER GUIDE. |  |  | Molecular Q2: 2; Molecular Q4: Answer all sub questions |  | 0 |  |  |  | 8.7
|  |  | Oxacillin If the laboratory uses an oxacillin disk (1ug) to screen for penicillin resistance (Penicillin G or Benzylpenicillin, the IV formulation) in S. pneumoniae and the zone size < 20, then the laboratory must do an MIC method before reporting penicillin as resistant (CLSI recommendation). EUCAST recommends that if the zone size is < 20mm to do a MIC, if ≧ 20 mm the result should be reported as susceptible. | [19 empty] | 100
|  |  | Co-trimoxazole | [19 empty] | 100
|  |  | Ceftriaxone or cefotaxime | [19 empty] | 100 | 300
GRAM NEGATIVE BACILLI ID & AST USING CONVENTIONAL METHODS
B8.16 |  | Is the following testing performed to identify gram negative bacilli? |  |  | Molecular Q2: 2; Molecular Q4: Answer all sub questions |  | 0 |  |  |  | 8.7
|  |  | Oxidase | [19 empty] | 100
|  |  | Indole | [19 empty] | 100
|  |  | Methyl Red | [19 empty] | 100
|  |  | Voges Proskauer | [19 empty] | 100
|  |  | Citrate | [19 empty] | 100
|  |  | Triple Sugar Iron or Kligler Iron | [19 empty] | 100
|  |  | Urease | [19 empty] | 100
|  |  | Motility | [19 empty] | 100 | 800
B8.17 |  | Does the lab follow the latest CLSI /EUCAST guidelines for AST of Gram negative bacilli? https://www.clsi.org / www.eucast.org NOTE: TECHNICAL INFORMATION RELATED TO THIS QUESTION IS PROVIDED IN THE USER GUIDE. |  |  | Molecular Q2: 2 |  | 0 |  |  |  | 8.7
B8.18 |  | Does the laboratory use Combination Disk Test or another equivalent method for Extended Spectrum Beta-Lactamase (ESBL) screening? (Reference: J Clin Microbiol. 2013 Sep; 51(9): 2986–2990. DOI: 10.1128/JCM.00901-13). NOTE: TECHNICAL INFORMATION RELATED TO THIS QUESTION IS PROVIDED IN THE USER GUIDE. |  |  | Molecular Q2: 2 |  | 0 |  |  |  | 8.7
B8.19 |  | Does the laboratory use Combination Disk Test or another equivalent method for carbapenemase screening? NOTE: TECHNICAL INFORMATION RELATED TO THIS QUESTION IS PROVIDED IN THE USER GUIDE. |  |  | Molecular Q2: 2 |  | 0 |  |  |  | 8.7
INTERLABORATORY COMPARISON, PT AND EXTERNAL QUALITY ASSURANCE (EQA) NOTE: TECHNICAL INFORMATION RELATED TO THE QUESTIONS IN THIS SECTION IS PROVIDED IN THE USER GUIDE.
B8.20 |  | Is the laboratory enrolled in an interlaboratory comparison or PT program for blood culture and molecular tests for organism identification, and AST? |  |  | Molecular Q2: 2 |  | 0 |  |  |  | 8.14
B8.21 |  | Did the laboratory pass the last 3 rounds of interlaboratory comparison or PT program testing? |  |  | Molecular Q2: 2 |  | 0 |  |  |  | 8.14
B8.22 |  | Does the laboratory receive onsite supervision visits as part of the EQA program for blood culture and molecular tests? |  |  | Molecular Q2: 2 |  | 0 |  |  |  | 8.14
Sub total |  |  |  |  | Molecular Q2: 45 |  | 0
Section 9: Information Management
All generic requirements apply, see SLIPTA Section 9. In addition, assessors should review the following:
Question no. |  | Question |  |  | Molecular Q2: Available score; Molecular Q4: Answer |  | Score |  | Comments |  | Related to SLIPTA question
B9.1 |  | Does the final report for blood culture list the organisms for which the specimen was and was not cultured? The laboratory should inform the clinician on the report what organisms were excluded during the culture process. This may be either by choice of media or incubation conditions (e.g. anaerobic organisms). Assessors should review a number of laboratory reports to determine how results are reported. Procedures should be consistent with the laboratory’s SOPs. |  |  | Molecular Q2: 2 |  | 0 |  |  |  | 9.3
B9.2 |  | Does the laboratory report alert organisms which include at least? (Alert organisms are organisms with significant public health threat and / or organisms that are notifiable). |  |  | Molecular Q2: 2; Molecular Q4: Answer all sub questions |  | 0 |  |  |  | 9.3
|  |  | Methicillin resistant S. aureus | [19 empty] | 100
|  |  | Imipenem restistant K. pneumoniae | [19 empty] | 100
|  |  | Carbapenem resistant Enterobacteriaceae | [19 empty] | 100
|  |  | ESBL producing organisms | [19 empty] | 100
|  |  | Multidrug resistant Pseudomonas | [19 empty] | 100
|  |  | Multidrug resistant Acinetobacter | [19 empty] | 100 | 600
Sub total |  |  |  |  | Molecular Q2: 4 |  | 0
Section 10: Identification of Nonconformities, Corrective and Preventive Actions
All generic requirements apply, see SLIPTA Section 10.
Section 11: Occurrence/Incident Management & Process Improvement
All generic requirements apply, see SLIPTA Section 11. In addition, assessors should review the following:
Question no. |  | Question |  |  | Molecular Q2: Available score; Molecular Q4: Answer |  | Score |  | Comments |  | Related to SLIPTA question
B11.1 |  | Are the following performance indicators collected? It may not be possible for laboratories to distinguish between community and hospital acquired infection if this is not collected on the laboratory requisition form. |  |  | Molecular Q2: 3; Molecular Q4: Answer all sub questions |  | 0 |  |  |  | 11.4 & 11.5
|  |  | Number of blood culture and molecular tests performed (disaggregated by type): |  | Molecular Q4: Answer all sub questions | [6 empty] | 100
|  |  |  | Hospital-acquired (bacterial infections in hospitalized patients (i.e. pathogenic bacterial isolated from a sample collected more than 48 hours after admission). | [18 empty] | 100
|  |  |  | Community-acquired (ambulatory patients and hospitalized patients from which a sample was collected less than 48 hours after admission). | [18 empty] | 100
|  |  |  | Unknown / referred If the laboratory can’t distinguish between hospital & community acquired infections, the number of organisms isolated should be recorded as "Unknown/referred". | [18 empty] | 100 | 300
|  |  | Number of blood culture and molecular tests where pathogens were isolated (disaggregated by type): |  | Molecular Q4: Answer all sub questions | [6 empty] | 100
|  |  |  | S. aureus | [18 empty] | 100
|  |  |  | S. pneumoniae | [18 empty] | 100
|  |  |  | K. pneumoniae | [18 empty] | 100
|  |  |  | A. baumannii | [18 empty] | 100
|  |  |  | E. coli | [18 empty] | 100
|  |  |  | Salmonella sp. | [18 empty] | 100 | 600
|  |  | Number and percentage of contaminated blood culture tests (disaggregated by in-patient & out-patient & unknown/referred) | [19 empty] | 100
|  |  | Bacterial blood culture and molecular tests TAT (disaggregated by in-patient/out-patient/unknown/referred - from sample collection to reporting) | [19 empty] | 100 | 400
Sub total |  |  |  |  | Molecular Q2: 3 |  | 0
Section 12: Facilities and Biosafety
All generic requirements apply, see SLIPTA Section 12. In addition, assessors should review the following:
Question no. |  | Question |  |  | Molecular Q2: Available score; Molecular Q4: Answer |  | Score |  | Comments |  | Related to SLIPTA question
B12.1 |  | Is a biological safety cabinet (BSC) or hood available and used for handling specimens or organisms considered to be highly contagious by airborne routes? |  |  | Molecular Q2: 2 |  | 0 |  |  |  | 12.8
B12.2 |  | Post-exposure prophylaxis: |  |  | Molecular Q2: 2; Molecular Q4: Answer all sub questions |  | 0 |  |  |  | 12.18
|  |  | a) Does the laboratory have a policy for cases of needlestick injury? | [19 empty] | 100
|  |  | b) Are anti-retroviral drugs (ARV) available for post-exposure prophylaxis (PEP) in case of needlestick injury and, if yes, are the drugs in date? | [19 empty] | 100 | 200
Sub total |  |  |  |  | Molecular Q2: 4 |  | 0
Totals
[5 empty] | Molecular Q2: 76 |  | 0
Links to other AMR modules | bar: Manual Q1
- General AMR module
- Urine module
- Feces module
- Genital module
- Pulmonary module
- Wound module
Yes
No
Yes
Partial
No
Yes
No
N/A
Yes
Partial
No
N/A

## Genital Module
AMR Lab Scorecard v9.2 (Sep 2021)
Bacterial culture, detection, identification and antimicrobial susceptibility testing of genital samples
N-A | How many genital culture and other tests were performed last year? | [15 empty] | <------- | It is highly recommended that assessors obtain the necessary permission to review the laboratory data. However, if assessors are unable to review the laboratory this question is NOT compulsory for completion of the assessment.
[5 empty] | Culture |  |  |  | Molecular |  |  |  | Other |  |  |  | <------- | Molecular = Molecular tests performed on genital samples for the detectoin of N. gonorrhoeae.
[5 empty] | Q1 | Q2 | Q3 | Q4 | Q1 | Q2 | Q3 | Q4 | Q1 | Q2 | Q3 | Q4
Adult vs. congenital
Male adult
Female adult
Congenital (children)
| TOTAL |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Symptomatic vs. contact tracing
Symptomatic
Contacts
| TOTAL |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Q = Quarter
N-B | Are there any significant variations (> 20%) in the number of cultures or other tests performed or organisms isolated each quarter? If ‘Yes’, please explain.
[14 empty] | <------- | It is highly recommended that assessors obtain the necessary permission to review the laboratory data. However, if assessors are unable to review the laboratory this question is NOT compulsory for completion of the assessment.
Section 1: Documents and records
All generic requirements apply, see SLIPTA Section 1. In addition, assessors should review the following:
Question no. |  | Question |  |  | Other Q2: Available score; Other Q4: Answer |  | Score | Comments | Related to SLIPTA question
N1.1 |  | Does the laboratory have documentation covering the following processes? |  |  | Other Q2: 2; Other Q4: Answer all sub questions |  | 0 |  | 1.5
|  |  | a) Production of selective agar (Thayer-Martin, Martin-Lewis, GC-lect® or similar) and non-selective agar (chocolate agar or similar) for N. gonorrhoeae isolation? | [17 empty] | 100
|  |  | b) Microscopic examination of smears | [17 empty] | 100
|  |  | c) Processing of samples for culture and molecular tests | [17 empty] | 100
|  |  | d) Detection, identification and AST of N. gonorrhoeae | [17 empty] | 100
|  |  | e) Reporting of N. gonorrhoeae culture and molecular test results | [17 empty] | 100
|  |  | f) Interlaboratory comparison or proficiency testing (PT) | [17 empty] | 100
|  |  | g) Laboratory safety | [17 empty] | 100 | 700
N1.2 |  | Are the documents complete, in-date and witnessed by all staff performing N. gonorrhoeae culture and molecular tests? |  |  | Other Q2: 2 |  | 0 |  | 1.5
See ISO15189:2012 Clause 5.5.3 for minimum requirements for a technical Standard Operating Procedure (SOP).
N1.3 |  | Are/is the following processes documented? |  |  | Other Q2: 3; Other Q4: Answer all sub questions |  | 0 |  | 1.5
|  |  | a) How to identify N. gonorrhoeae on all primary media? (SOP should describe colony appearance of N. gonorrhoeae and define how to proceed when potential N. gonorrhoeae is encountered) | [17 empty] | 100
|  |  | b) Instructions for handling samples received after hours? | [17 empty] | 100
|  |  | c) Instructions for referral of bacterial isolates for identification and AST? | [17 empty] | 100
|  |  | d) Instructions on how to perform AST conversions for automated, disk diffusion, Etest / Gradient and microdilution AST? | [17 empty] | 100
|  |  | e) Turnaround time for N. gonorrhoeae culture or molecular tests? From sample collection to reporting. | [17 empty] | 100 | 500
Sub total |  |  |  |  | Other Q2: 7 |  | 0
Section 2: Management Reviews
All generic requirements apply, see SLIPTA Section 2.
Section 3: Organization & Personnel
All generic requirements apply, see SLIPTA Section 3. In addition, assessors should review the following:
Question no. |  | Question |  |  | Other Q2: Available score; Other Q4: Answer |  | Score | Comments | Related to SLIPTA question
N3.1 |  | Is there evidence that laboratory staff have been trained and are competent in the following: |  |  | Other Q2: 3; Other Q4: Answer all sub questions |  | 0 |  | 3.6
|  |  | a) Processing of samples for N. gonorrhoeae culture and molecular tests | [17 empty] | 100
|  |  | b) Detection / identification and AST of N. gonorrhoeae | [17 empty] | 100
|  |  | c) Interpretation of microscopy smears for the detection of N. gonorrhoeae | [17 empty] | 100
|  |  | d) Interpretation of N. gonorrhoeae culture and molecular test results, taking into account any clinical information | [17 empty] | 100
|  |  | e) Reporting of N. gonorrhoeae culture and molecular test results | [17 empty] | 100
|  |  | f) QC, EQA & PT for N. gonorrhoeae culture and molecular tests | [17 empty] | 100
|  |  | g) Laboratory safety | [17 empty] | 100 | 700
Review training records, competency assessment forms and duty rosters. Pay attention to date of training and scope of training compared with techniques being performed.
N3.2 |  | Is there evidence that laboratory staff are following the procedures described in the laboratory documentation? |  |  | Other Q2: 3; Other Q4: Answer all sub questions |  | 0 |  | 3.7
|  |  | a) Processing of samples for N. gonorrhoeae culture and molecular tests | [17 empty] | 100
|  |  | b) Interpretation of microscopy smears for the detection of N. gonorrhoeae | [17 empty] | 100
|  |  | c) Interpretation of N. gonorrhoeae culture and molecular test results | [17 empty] | 100
|  |  | d) Identification and AST of N. gonorrhoeae | [17 empty] | 100
|  |  | e) Reporting of N. gonorrhoeae culture and molecular test results | [17 empty] | 100 | 500
Directly observe procedures being performed compared to the SOP.
Sub total |  |  |  |  | Other Q2: 6 |  | 0
Section 4: Client Management & Customer Service
All generic requirements apply, see SLIPTA Section 4. In addition, assessors should review the following:
Question no. |  | Question |  |  | Other Q2: Available score; Other Q4: Answer |  | Score | Comments | Related to SLIPTA question
N4.1 |  | Is there evidence that the laboratory has provided clients information / instructions on N. gonorrhoeae sample collection, storage and transportation to the laboratory? |  |  | Other Q2: 3 |  | 0 |  | 4.1
N4.2 |  | Is there evidence that the laboratory has provided clients information / instructions on interpretation of N. gonorrhoeae culture test results and AST? |  |  | Other Q2: 2 |  | 0 |  | 4.1
Sub total |  |  |  |  | Other Q2: 5 |  | 0
Section 5: Equipment
All generic requirements apply, see SLIPTA Section 5.
Section 6: Evaluation and Audits
All generic requirements apply, see SLIPTA Section 6. Note: this checklist (with or without SLIPTA) can be used for internal and external audits.
Section 7: Purchasing & Inventory
All generic requirements apply, see SLIPTA Section 7. In addition, assessors should review the following:
Question no. |  | Question |  |  | Other Q2: Available score; Other Q4: Answer |  | Score | Comments | Related to SLIPTA question
N7.1 |  | Is all media for N. gonorrhoeae culture, isolation, identification and AST stored correctly and in date (from date of manufacture media must be stored at 2-8°C)? According to manufacturer’s requirements. |  |  | Other Q2: 2; Other Q4: Answer all sub questions |  | 0 |  | 7.10
|  |  | - Selective media | [17 empty] | 100
|  |  | - Non-selective media | [17 empty] | 100 | 200
Sub total |  |  |  |  | Other Q2: 2 |  | 0
Section 8: Process Control
All generic requirements apply, see SLIPTA Section 8. In addition, assessors should review the following:
Question no. |  | Question |  |  | Other Q2: Available score; Other Q4: Answer |  | Score | Comments | Related to SLIPTA question
SPECIMEN COLLECTION
N8.1 |  | Does the laboratory request form have space for indicating gender, age, whether treatment has started (incl. start date) and whether the case is symptomatic or asymptomatic (contact tracing)? |  |  | Other Q2: 2 |  | 0 |  | 4.1
N8.2 |  | Are two samples collected (one for culture and one for PCR)? |  |  | Other Q2: 2 |  | 0 |  | 4.1
N8.3 |  | If multiple specimens are collected, are specimens for N. gonorrhoeae culture collected first? |  |  | Other Q2: 2 |  | 0 |  | 4.1
N8.4 |  | Has the laboratory ensured use of correct sample transportation systems (incl. transportation media)? |  |  | Other Q2: 2 |  | 0 |  | 4.1
N8.5 |  | Has the laboratory made sure that sample transportation time is kept under 12 hours? |  |  | Other Q2: 2 |  | 0 |  | 4.1
N8.6 |  | Does the laboratory monitor sample rejection rate and take action when trends point to problems related to sample collection and transportation? |  |  | Other Q2: 2 |  | 0 |  | 4.1
MEDIA QUALITY CONTROL
N8.7 |  | Does the laboratory perform QC testing on all media before use? This includes in-house made or purchased from commercial sources. |  |  | Other Q2: 3; Other Q4: Answer all sub questions |  | 0 |  | 8.8
Selective media (Thayer-Martin/Martin-Lewis/GC-Lect® or equivalent)
|  |  | Do QC records demonstrate that they are checked for their ability to suppress growth of normal flora (i.e. for anogential and/or nasopharyngyal samples) while allowing the growth of N. gonorrhoeae? | [17 empty] | 100
Non-selective media (GC-chocolate agar or equivalent)
|  |  | Do QC records demonstrate that the nonselective media is checked for their ability to support growth of N. gonorrhoeae? | [17 empty] | 100 | 200
N8.8 |  | Does the laboratory: NOTE: TECHNICAL INFORMATION RELATED TO THIS QUESTION IS PROVIDED IN THE USER GUIDE. |  |  | Other Q2: 2; Other Q4: Answer all sub questions |  | 0 |  | 8.8
|  |  | a) Perform sterility and performance tests for every batch of culture media using certified reference strains as controls? | [17 empty] | 100
|  |  | b) Are reference strains sourced from an authorized supplier (e.g. ATCC)? | [17 empty] | 100
|  |  | c) Are the reference strains stored, cultured and sub-cultured in accordance with the specification from the supplier? | [17 empty] | 100 | 300
N8.9 |  | Does the laboratory determine the cause of failed QC (root cause analysis), perform corrective actions and measure the effectiveness thereof? |  |  | Other Q2: 2 |  | 0 |  | 8.10
BACTERIAL GENITAL CULTURE PROCEDURE
N8.10 |  | Are the following media used for primary culture of genital samples? Either Modified Thayer-Martin (MTM) agar, Martin-Lewis (ML) agar, GC-Lect® or equivalent, and GC-chocolate agar or equivalent. |  |  | Other Q2: 2; Other Q4: Answer all sub questions |  | 0 |  | 8.7
|  |  | MTM agar, ML agar, GC-Lect® or equivalent | [17 empty] | 100
|  |  | GC-chocolate agar or equivalent | [17 empty] | 100 | 200
N8.11 |  | Are media used for primary culture of N. gonorrhoeae incubated at 35-37°C for at least 24 hours? |  |  | Other Q2: 2 |  | 0 |  | 8.7
N8.12 |  | Are media used for primary culture of N. gonorrhoeae incubated aerobically, in a humid environment with 5% CO2? Use a CO2 incubator or an alternative method. |  |  | Other Q2: 2 |  | 0 |  | 8.7
N. GONORRHOEAE IDENTIFICATION AND AST BY CONVENTIONAL METHODS NOTE: TECHNICAL INFORMATION RELATED TO THE QUESTIONS IN THIS SECTION IS PROVIDED IN THE USER GUIDE.
N8.13 |  | For automated, kit-based, molecular or conventional methods for N.gonorrhoeae only: |  |  | Other Q2: 2 |  | 0 |  | 8.7
When performing AST, are isolates used that have been incubated between 18 and 24 hours?
N8.14 |  | Is identity of N. gonorrhoeae confirmed with biochemical tests to rule-out commensal Neisseria or N. meningitidis? |  |  | Other Q2: 2 |  | 0 |  | 8.7
N8.15 |  | Is the following testing performed for N. gonorrhoeae identification: |  |  | Other Q2: 2; Other Q4: Answer all sub questions |  | 0 |  | 8.7
|  |  | - Gram stain | [17 empty] | 100
|  |  | - Oxidase | [17 empty] | 100
|  |  | - Catalase | [17 empty] | 100
|  |  | - Immunological tests | [17 empty] | 100
|  |  | - Biochemical identification tests/kits | [17 empty] | 100 | 500
N8.16 |  | Is N. gonorrhoeae AST done as per current WHO or other approved guidelines? NOTE: TECHNICAL INFORMATION RELATED TO THIS QUESTION IS PROVIDED IN THE USER GUIDE. |  |  | Other Q2: 2 |  | 0 |  | 8.7
INTERLABORATORY COMPARISON, PT AND EXTERNAL QUALITY ASSESSMENT (EQA) NOTE: TECHNICAL INFORMATION RELATED TO THE QUESTIONS IN THIS SECTION IS PROVIDED IN THE USER GUIDE.
N8.17 |  | Is the laboratory enrolled in an interlaboratory comparison or PT program for N. gonorrhoeae culture and molecular tests for organism identification and AST? |  |  | Other Q2: 2 |  | 0 |  | 8.14
N8.18 |  | Did the laboratory pass the last 3 rounds of interlaboratory comparison or PT program testing? |  |  | Other Q2: 2 |  | 0 |  | 8.14
N8.19 |  | Does the laboratory receive onsite supervision visits as part of the EQA program for N. gonorrhoeae culture and molecular tests? |  |  | Other Q2: 2 |  | 0 |  | 8.14
Sub total |  |  |  |  | Other Q2: 39 |  | 0
Section 9: Information Management
All generic requirements apply, see SLIPTA Section 9. In addition, assessors should review the following:
Question no. |  | Question |  |  | Other Q2: Available score; Other Q4: Answer |  | Score | Comments | Related to SLIPTA question
N9.1 |  | Does the final report for N. gonorrhoeae culture list all the antibiotics for which the specimen was tested? The laboratory should inform the clinician for which antibiotics susceptibility was determined. Antibiotics to be tested for N. gonorrhoeae include 3rd generation cephalosporins (Cefixime and Ceftriaxone), fluoroquinolones (Ciprofloxacin), macrolides (Azithromycin), aminocyclitols (Spectinomycin), and aminoglycosides (Gentamycin). |  |  | Other Q2: 2 |  | 0 |  | 9.3
Sub total |  |  |  |  | Other Q2: 2 |  | 0
Section 10: Identification of Nonconformities, Corrective and Preventive Actions
All generic requirements apply, see SLIPTA Section 10.
Section 11: Occurrence/Incident Management & Process Improvement
All generic requirements apply, see SLIPTA Section 11. In addition, assessors should review the following:
Question no. |  | Question |  |  | Other Q2: Available score; Other Q4: Answer |  | Score | Comments | Related to SLIPTA question
N11.1 |  | Are the following performance indicators collected? |  |  | Other Q2: 3; Other Q4: Answer all sub questions |  | 0 |  | 11.4 & 11.5
|  |  | Number of N. gonorrhoeae culture and molecular tests performed (disaggregated by type): |  | Other Q4: Answer all sub questions |  |  |  |  | 100
|  |  |  | Male vs. female vs. congenital (children) | [16 empty] | 100
|  |  |  | Symptomatic cases vs. contacts | [16 empty] | 100
|  |  |  | By sample type | [16 empty] | 100 | 300
|  |  | Number and percentage of genital samples rejected (disaggregated by reason e.g. leaked, transportation time too long) (target <1%) | [17 empty] | 100
|  |  | Number of culture tests where N. gonorrhoeae was isolated or identified | [17 empty] | 100
|  |  | N. gonorrhoeae culture and molecular tests TAT (disaggregated by in-patient/out-patient/unknown/referred - from sample collection to reporting) | [17 empty] | 100 | 400
Sub total |  |  |  |  | Other Q2: 3 |  | 0
Section 12: Facilities and Biosafety
All generic requirements apply, see SLIPTA Section 12.
Totals
[5 empty] | Other Q2: 64 |  | 0
Links to other AMR modules | bar: Culture Q1
- General AMR module
- Urine module
- Feces module
- Blood module
- Pulmonary module
- Wound module
Yes
No
Yes
Partial
No
N/A
Yes
Partial
No
Yes
No
N/A

## Pulmonary Module
AMR Lab Scorecard v9.2 (Sep 2021)
Bacterial culture, detection, identification and antimicrobial susceptibility testing of pulmonary samples
P-A | How many pulmonary sample culture tests and molecular tests were performed last year? | [12 empty] | <------- | It is highly recommended that assessors obtain the necessary permission to review the laboratory data. However, if assessors are unable to review the laboratory this question is NOT compulsory for completion of the assessment.
[5 empty] | Culture |  |  |  | Molecular
[5 empty] | Q1 | Q2 | Q3 | Q4 | Q1 | Q2 | Q3 | Q4 | Sub total | <------- | See: http://www.who.int/glass/en/ and other frequently isolated pathogens.
| Hospital-acquired | [12 empty] | <------- | Hospital-acquired infections are defined as bacterial infections in hospitalized patients (i.e. pathogenic bacterial isolated from a sample collected more than 48 hours after admission).
|  | S. aureus | [10 empty] | 0
|  | S. pneumoniae | [10 empty] | 0
|  | S. pyogenes | [10 empty] | 0
|  | Moraxella catarrhalis | [10 empty] | 0
|  | C. diphteriae | [10 empty] | 0
|  | H. influenzae | [10 empty] | 0
|  | K. pneumoniae | [10 empty] | 0
|  | Mycoplasma pneumoniae | [10 empty] | 0
| Community-acquired | [12 empty] | <------- | Community-acquired infections are defined as ambulatory patients and hospitalized patients from which a sample was collected less than 48 hours after admission.
|  | S. aureus | [10 empty] | 0
|  | S. pneumoniae | [10 empty] | 0
|  | S. pyogenes | [10 empty] | 0
|  | Moraxella catarrhalis | [10 empty] | 0
|  | C. diphteriae | [10 empty] | 0
|  | H. influenzae | [10 empty] | 0
|  | K. pneumoniae | [10 empty] | 0
|  | Mycoplasma pneumoniae | [10 empty] | 0
| Unknown/referred | [12 empty] | <------- | If the laboratory can’t distinguish between hospital & community acquired infections, the number of organisms isolated should be recorded as "Unknown/referred".
|  | S. aureus | [10 empty] | 0
|  | S. pneumoniae | [10 empty] | 0
|  | S. pyogenes | [10 empty] | 0 |  | Total number of isolates per pathogen
|  | Moraxella catarrhalis | [10 empty] | 0
|  | C. diphteriae | [10 empty] | 0 |  | S. aureus |  | 0
|  | H. influenzae | [10 empty] | 0 |  | S. pneumoniae |  | 0
|  | K. pneumoniae | [10 empty] | 0 |  | S. pyogenes |  | 0
|  | Mycoplasma pneumoniae | [10 empty] | 0 |  | Moraxella catarrhalis |  | 0
| TOTAL ISOLATES |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | C. diphteriae |  | 0
| TOTAL NUMBER OF PULMONARY TESTS PERFORMED | [13 empty] | H. influenzae |  | 0
| TOTAL NUMBER OF PULMONARY SAMPLE CULTURES WITH NO PATHOGENS ISOLATED/IDENTIFIED | [13 empty] | K. pneumoniae |  | 0
[15 empty] | Mycoplasma pneumoniae |  | 0
Q = Quarter
P-B | Are there any significant variations (>20%) in the number of pulmonary sample culture tests perfromed or organisms isolated or identified each quarter? If 'yes', please explain.
[13 empty] | <------- | It is highly recommended that assessors obtain the necessary permission to review the laboratory data. However, if assessors are unable to review the laboratory this question is NOT compulsory for completion of the assessment.
Section 1: Documents and records
All generic requirements apply, see SLIPTA Section 1. In addition, assessors should review the following:
Question no. |  | Question | [10 empty] | Available score |  | Answer | Score | Comments | Related to SLIPTA question
P1.1 |  | Does the laboratory have documentation covering the following processes? | [10 empty] | 2 |  | Answer all sub questions | 0 |  | 1.5
|  |  | a) Production of Blood Agar, MacConkey Agar or other media for pulmonary pathogen isolation? | [15 empty] | 100
|  |  | b) Processing of pulmonary samples | [15 empty] | 100
|  |  | c) Detection, identification and AST of pulmonary pathogens | [15 empty] | 100
|  |  | d) Reporting of pulmonary sample culture and molecular test results | [15 empty] | 100
|  |  | e) Interlaboratory comparison or proficiency testing (PT) | [15 empty] | 100
|  |  | f) Laboratory safety | [15 empty] | 100 | 600
P1.2 |  | Are the documents complete, in-date and witnessed by all staff performing N. gonorrhoeae culture and molecular tests? | [10 empty] | 2 |  |  | 0 |  | 1.5
See ISO15189:2012 Clause 5.5.3 for minimum requirements for a technical Standard Operating Procedure (SOP).
P1.3 |  | Are/is the following processes documented? | [10 empty] | 3 |  | Answer all sub questions | 0 |  | 1.5
|  |  | a) Rejection criteria for pulmonary samples? | [15 empty] | 100
|  |  | b) A policy for reporting critical pulmonary results? | [15 empty] | 100
|  |  | c) Instructions for reporting pulmonary sample culture tests with mixed bacterial growth? | [15 empty] | 100
|  |  | d) Instructions for referral of pulmonary sample culture or molecular tests at the laboratory? | [15 empty] | 100
|  |  | e) Instructions for handling samples received after hours? | [15 empty] | 100
|  |  | f) Instructions for referral of bacterial isolates for identification and AST? | [15 empty] | 100
|  |  | g) Instructions on how to perform AST conversions for automated, disk diffusion, Etest / Gradient and microdilution AST? | [15 empty] | 100
|  |  | h) Turnaround time for pulmonary sample culture or molecular tests? From sample collection to reporting. | [15 empty] | 100
|  |  | i) Definition of rare / unexpected AST results? | [15 empty] | 100
|  |  | j) Confirmatory tests for unusual or unexpected patient AST results? | [15 empty] | 100 | 1000
Sub total | [12 empty] | 7 |  |  | 0
Section 2: Management Reviews
All generic requirements apply, see SLIPTA Section 2.
Section 3: Organization & Personnel
All generic requirements apply, see SLIPTA Section 3. In addition, assessors should review the following:
Question no. |  | Question | [10 empty] | Available score |  | Answer | Score | Comments | Related to SLIPTA question
P3.1 |  | Is there evidence that laboratory staff have been trained and are competent in the following: | [10 empty] | 3 |  | Answer all sub questions | 0 |  | 3.6
|  |  | a) Processing of pulmonary samples for culture and molecular tests | [15 empty] | 100
|  |  | b) Identification and AST of pulmonary pathogens | [15 empty] | 100
|  |  | c) Interpretation of pulmonary sample culture and molecular test results | [15 empty] | 100
|  |  | d) Reporting of pulmonary sample culture and molecular test results | [15 empty] | 100
|  |  | e) QC, EQA & PT for pulmonary sample culture and molecular tests | [15 empty] | 100
|  |  | f) Laboratory safety | [15 empty] | 100 | 600
Review training records, competency assessment forms and duty rosters. Pay attention to date of training and scope of training compared with techniques being performed.
P3.2 |  | Is there evidence that laboratory staff are following the procedures described in the laboratory documentation? | [10 empty] | 3 |  | Answer all sub questions | 0 |  | 3.7
|  |  | a) Processing of pulmonary for culture and molecular tests | [15 empty] | 100
|  |  | b) Interpretation of pulmonary sample culture test results | [15 empty] | 100
|  |  | c) Identification and AST of pulmonary pathogens | [15 empty] | 100
|  |  | d) Reporting of pulmonary sample culture test and molecular test results | [15 empty] | 100 | 400
Directly observe procedures being performed compared to the SOP.
Sub total | [12 empty] | 6 |  |  | 0
Section 4: Client Management & Customer Service
All generic requirements apply, see SLIPTA Section 4. In addition, assessors should review the following:
Question no. |  | Question | [10 empty] | Available score |  | Answer | Score | Comments | Related to SLIPTA question
P4.1 |  | Please answer the questions below: | [10 empty] | 3 |  | Answer all sub questions |  |  | 4.1
|  |  | a) Is there evidence that the laboratory has provided clients information / instructions on pulmonary sample collection, storage and transportation to the laboratory? | [15 empty] | 100
|  |  | b) Does the information include selection of appropriate type of specimen? | [15 empty] | 100 | 200
P4.2 |  | Is there evidence that the laboratory has provided clients information / instructions on interpretation of pulmonary sample culture results and AST? | [10 empty] | 2 |  |  | 0 |  | 4.1
Sub total | [12 empty] | 5 |  |  | 0
Section 5: Equipment
All generic requirements apply, see SLIPTA Section 5.
Section 6: Evaluation and Audits
All generic requirements apply, see SLIPTA Section 6. Note: this checklist (with or without SLIPTA) can be used for internal and external audits.
Section 7: Purchasing & Inventory
All generic requirements apply, see SLIPTA Section 7. In addition, assessors should review the following:
Question no. |  | Question | [10 empty] | Available score |  | Answer | Score | Comments | Related to SLIPTA question
P7.1 |  | Is all media for bacterial culture isolation, identification and AST stored correctly and in date (from date of manufacture media must be stored at 2-8°C)? According to manufacturer's requirements. | [10 empty] | 2 |  | Answer all sub questions | 0 |  | 7.10
|  |  | - Blood agar | [15 empty] | 100
|  |  | - MacConkey agar | [15 empty] | 100
|  |  | - Chocolate agar | [15 empty] | 100
|  |  | - Tellurite agar (or equivalent) | [15 empty] | 100
|  |  | - New York City medium (or equivalent) | [15 empty] | 100
|  |  | - Mueller Hinton | [15 empty] | 100 | 600
Sub total | [12 empty] | 2 |  |  | 0
Section 8: Process Control
All generic requirements apply, see SLIPTA Section 8. In addition, assessors should review the following:
Question no. |  | Question | [10 empty] | Available score |  | Answer | Score | Comments | Related to SLIPTA question
MEDIA QUALITY CONTROL
P8.1 |  | Does the laboratory perform QC testing on all media before use? This includes in-house made or purchased from commercial sources. | [10 empty] | 3 |  | Answer all sub questions | 0 |  | 8.8
Blood agar
|  |  | Do QC records for blood agar plates demonstrate that they are checked for their ability to support growth of fastidious organisms such as S. pneumoniae? | [15 empty] | 100
|  |  | Do QC records for blood agar plates demonstrate that they are checked for their ability to show beta, alpha, and gamma hemolysis? | [15 empty] | 100
MacConkey agar (MAC)
|  |  | Do QC records for MAC plates demonstrate that they are checked for their ability to suppress growth of Gram-positive organisms while allowing the growth of Gram-negative organisms? | [15 empty] | 100
|  |  | Do QC records for MAC plates demonstrate that they are checked for their ability to allow visualization of lactose fermentation? | [15 empty] | 100
Chocolate agar
|  |  | Do QC records for Chocolate agar plates demonstrate that they are checked for their ability to support growth of fastidious organisms such as H. influenzae? | [15 empty] | 100
Tellurite agar (or equivalent)
|  |  | Do QC records for Tellurite agar plates demonstrate that they are checked for their ability to support growth of fastidious organisms such as C. diphtheriae? | [15 empty] | 100
New York City medium (or equivalent)
|  |  | Do QC records for New York City medium (or equivalent) demonstrate that they are checked for their ability to support growth of fastidious organisms such as M. pneumoniae? | [15 empty] | 100
Mueller Hinton Agar (MHA)
|  |  | Do QC records demonstrate that MHA plates are checked for their ability to grow S. aureus & E. coli? | [15 empty] | 100 | 800
P8.2 |  | Does the laboratory: | [10 empty] | 2 |  | Answer all sub questions | 0 |  | 8.8
|  |  | a) Perform sterility and performance tests for every batch of culture media using certified reference strains as controls? | [15 empty] | 100
|  |  | b) Are reference strains sourced from an authorized supplier (e.g. ATCC)? | [15 empty] | 100
|  |  | c) Are the reference strains stored, cultured and sub-cultured in accordance with the specification from the supplier? | [15 empty] | 100 | 300
P8.3 |  | Does the laboratory determine the cause of failed QC (root cause analysis), perform corrective actions and measure the effectiveness thereof? | [10 empty] | 2 |  |  | 0 |  | 8.10 | 100
PULMONARY SAMPLE CULTURE PROCEDURE
For complete recommended procedure, see the User Guide.
P8.4 |  | Are pulmonary sample cultures plated onto selective and non-selective media including (at least): Media used for primary isolation may be adapted for sample type, see the User Guide. | [10 empty] | 2 |  | Answer all sub questions | 0 |  | 8.7
|  |  | Blood agar | [15 empty] | 100
|  |  | MacConkey agar | [15 empty] | 100
|  |  | Chocolate agar | [15 empty] | 100
|  |  | Tellurite agar (or equivalent) | [15 empty] | 100
|  |  | New York City medium (or equivalent) | [15 empty] | 100 | 500
P8.5 |  | Are pulmonary sample culture plates incubated at 35-37 degrees Celsius? | [10 empty] | 2 |  |  | 0 |  | 8.7 | 100
P8.6 |  | Does the laboratory report pulmonary sample cultures as contaminated if they contain organisms that should be considered contaminants (e.g. Streptococcus viridans)? | [10 empty] | 2 |  |  | 0 |  | 8.7 | 100
BACTERIAL ID & AST
P8.7 |  | Is the following testing performed for S. aureus identification? | [10 empty] | 2 |  | Answer all sub questions | 0 |  | 8.7
|  |  | - Catalase | [15 empty] | 100
|  |  | - Coagulase (slide or tube) | [15 empty] | 100
|  |  | - Mannitol Salt Agar (MSA) | [15 empty] | 100
|  |  | - Dnase | [15 empty] | 100 | 400
P8.8 |  | Does S. aureus AST include the following antibiotics: If oxacillin and cefoxitin results are discrepant for S. aureus (one is susceptible and one is resistant), the laboratory should repeat the testing. Note: oxacillin testing should always be tested by MIC (not disc diffusion). If the results remain discrepant, oxacillin should be reported as resistant. | [10 empty] | 2 |  | Answer all sub questions | 0 |  | 8.7
|  |  | - Cefoxitin | [15 empty] | 100
|  |  | - Vancomycin | [15 empty] | 100 | 200
P8.9 |  | Does the laboratory detect methicillin/nafcillin resistance in S. aureus using oxacillin disk? | [10 empty] | 2 |  |  | 0 |  | 8.7 | 100
P8.10 |  | Is the following testing performed for Streptococcus sp. identification? | [10 empty] | 2 |  | Answer all sub questions | 0 |  | 8.7
|  |  | - Bacitracin | [15 empty] | 100
|  |  | - Pyrrolidonyl Arylamidase (PYR) | [15 empty] | 100
|  |  | - Bile solubility | [15 empty] | 100
|  |  | - Optochin | [15 empty] | 100
|  |  | - S. pneumoniae latex | [15 empty] | 100 | 500
P8.11 |  | Does Streptococcus sp. AST include the following antibiotics: | [10 empty] | 2 |  | Answer all sub questions | 0 |  | 8.7
|  |  | - Oxacillin If the laboratory uses an oxacillin disk (1ug) to screen for penicillin resistance (Penicillin G or Benzylpenicillin, the IV formulation) in S. pneumoniae and the zone size < 20, then the laboratory must do an MIC method before reporting penicillin as resistant (CLSI recommendation). EUCAST recommends that if the zone size is < 20mm to do a MIC, if ≧ 20 mm the result should be reported as susceptible. | [15 empty] | 100
|  |  | - Co-trimoxazole | [15 empty] | 100
|  |  | - Ceftriaxone or cefotaxime | [15 empty] | 100 | 300
P8.12 |  | Is the following testing performed to identify Gram negative bacilli? | [10 empty] | 2 |  | Answer all sub questions | 0 |  | 8.7
|  |  | - Oxidase | [15 empty] | 100
|  |  | - Indole | [15 empty] | 100
|  |  | - Methyl Red | [15 empty] | 100
|  |  | - Voges Proskauer | [15 empty] | 100
|  |  | - Citrate | [15 empty] | 100
|  |  | - Triple Sugar Iron or Kligler Iron | [15 empty] | 100
|  |  | - Urease | [15 empty] | 100
|  |  | - Motility | [15 empty] | 100 | 800
P8.13 |  | Is the following testing performed to identify Moraxella catarrhalis? | [10 empty] | 2 |  | Answer all sub questions | 0 |  | 8.7
|  |  | - Oxidase | [15 empty] | 100
|  |  | - Tributyrin (CATScreen) | [15 empty] | 100
|  |  | - Dnase | [15 empty] | 100 | 300
P8.14 |  | Does Moraxella catarrhalis AST include the following antibiotics: | [10 empty] | 2 |  | Answer all sub questions | 0 |  | 8.7
|  |  | - Amoxicillin/clavulanic acid | [15 empty] | 100
|  |  | - Ceftriaxone or cefotaxime | [15 empty] | 100 | 200
P8.15 |  | Is the following testing performed to identify C. diphtheriae: | [10 empty] | 2 |  | Answer all sub questions | 0 |  | 8.7
|  |  | - Cystinase | [15 empty] | 100
|  |  | - Pyrazinamidase | [15 empty] | 100 | 200
P8.16 |  | Does C. diphtheriae AST include the following antibiotics: | [10 empty] | 2 |  | Answer all sub questions | 0 |  | 8.7
|  |  | - Penicillin | [15 empty] | 100
|  |  | - Erythromycin | [15 empty] | 100 | 200
P8.17 |  | Is the following testing performed to identify H. influenzae: | [10 empty] | 2 |  | Answer all sub questions | 0 |  | 8.7
|  |  | - X and V factor | [15 empty] | 100
|  |  | - H. influenzae serotyping | [15 empty] | 100 | 200
P8.18 |  | Does H. influenzae AST include the following antibiotics: | [10 empty] | 2 |  | Answer all sub questions | 0 |  | 8.7
|  |  | - Amoxicillin | [15 empty] | 100
|  |  | - Ceftriaxone or cefotaxime | [15 empty] | 100 | 200
P8.19 |  | Does the lab follow the latest CLSI /EUCAST guidelines for AST of Gram negative bacilli? https://www.clsi.org / www.eucast.org/) | [10 empty] | 2 |  |  | 0 |  | 8.7 | 100
P8.20 |  | Does the laboratory use Combination Disk Test or another equivalent method for Extended Spectrum Beta-Lactamase (ESBL) screening? J Clin Microbiol. 2013 Sep; 51(9): 2986–2990. | [10 empty] | 2 |  |  | 0 |  | 8.7 | 100
P8.21 |  | Does the laboratory use Combination Disk Test or another equivalent method for carbapenemase screening? | [10 empty] | 2 |  |  | 0 |  | 8.7 | 100
INTERLABORATORY COMPARISON, PT AND EXTERNAL QUALITY ASSURANCE (EQA)
P8.22 |  | Is the laboratory enrolled in an interlaboratory comparison or PT program for pulmonary sample culture and molecular tests for organism identification, and AST? | [10 empty] | 2 |  |  | 0 |  | 8.14
P8.23 |  | Did the laboratory pass the last 3 rounds of interlaboratory comparison or PT program testing? | [10 empty] | 2 |  |  | 0 |  | 8.14
P8.24 |  | Does the laboratory receive onsite supervision visits as part of the EQA program for pulmonary sample culture and molecular tests? | [10 empty] | 2 |  |  | 0 |  | 8.14
Sub total | [12 empty] | 49 |  |  | 0
Section 9: Information Management
All generic requirements apply, see SLIPTA Section 9. In addition, assessors should review the following:
Question no. |  | Question | [10 empty] | Available score |  | Answer | Score | Comments | Related to SLIPTA question
P9.1 |  | Does the final report for pulmonary sample culture list the organisms for which the specimen was and was not cultured? The laboratory should inform the clinician on the report what organisms were excluded during the culture process. This may be either by choice of media or incubation conditions (e.g. anaerobic organisms). Assessors should review a number of laboratory reports to determine how results are reported. Procedures should be consistent with the laboratory’s SOPs. | [10 empty] | 2 |  |  | 0 |  | 9.3
P9.2 |  | Does the laboratory report alert organisms which include at least? (Alert organisms are organisms with significant public health threat and / or organisms that are notifiable). | [10 empty] | 2 |  | Answer all sub questions | 0 |  | 9.3
|  |  | Methicillin resistant S. aureus | [15 empty] | 100
|  |  | Imipenem resistant Enterobacteriaceae | [15 empty] | 100
|  |  | Carbapenem resistant Enterobacteriaceae | [15 empty] | 100
|  |  | ESBL producing organisms | [15 empty] | 100
|  |  | K. pneumoniae | [15 empty] | 100 | 500
Sub total | [12 empty] | 4 |  |  | 0
Section 10: Identification of Nonconformities, Corrective and Preventive Actions
All generic requirements apply, see SLIPTA Section 10.
Section 11: Occurrence/Incident Management & Process Improvement
All generic requirements apply, see SLIPTA Section 11. In addition, assessors should review the following:
Question no. |  | Question | [10 empty] | Available score |  | Answer | Score | Comments | Related to SLIPTA question
P11.1 |  | Are the following performance indicators collected? It may not be possible for laboratories to distinguish between community and hospital acquired infection if this is not collected on the laboratory requisition form. | [10 empty] | 3 |  | Answer all sub questions | 0 |  | 11.4 & 11.5
|  |  | Number of pulmonary sample culture and molecular tests performed (disaggregated by type) | [11 empty] | Answer all sub questions |  |  |  | 100
|  |  |  | Hospital-acquired Hospital-acquired infections are defined as bacterial infections in hospitalized patients (i.e. pathogenic bacterial isolated from a sample collected more than 48 hours after admission). | [14 empty] | 100
|  |  |  | Community-acquired Community-acquired infections are defined as ambulatory patients and hospitalized patients from which a sample was collected less than 48 hours after admission. | [14 empty] | 100
|  |  |  | Unknown/referred If the laboratory can’t distinguish between hospital & community acquired infections, the number of organisms isolated should be recorded as "Unknown/referred". | [14 empty] | 100 | 300
|  |  | Number of pulmonary sample culture and molecular tests where pathogens were isolated (disaggregated by type): | [11 empty] | Answer all sub questions |  |  |  | 100
|  |  |  | S. aureus | [14 empty] | 100
|  |  |  | S. pneumoniae | [14 empty] | 100
|  |  |  | S. pyogenes | [14 empty] | 100
|  |  |  | Moraxella catarrhalis | [14 empty] | 100
|  |  |  | C. diphteria | [14 empty] | 100
|  |  |  | H. influenza | [14 empty] | 100
|  |  |  | K. pneumoniae | [14 empty] | 100
|  |  |  | Mycoplasma pneumoniae | [14 empty] | 100 | 800
|  |  | Pulmonary sample culture and molecular test TAT (disaggregated by in-patient & out-patient and by type - from sample collection to reporting) | [15 empty] | 100 | 300
Sub total | [12 empty] | 3 |  |  | 0
Section 12: Facilities and Biosafety
All generic requirements apply, see SLIPTA Section 12. In addition, assessors should review the following:
Question no. |  | Question | [10 empty] | Available score |  | Answer | Score | Comments | Related to SLIPTA question
P12.1 |  | Is a biological safety cabinet (BSC) or hood available and used for handling specimens or organisms considered to be highly contagious by air borne routes? | [10 empty] | 2 |  |  | 0 |  | 12.8
Sub total | [12 empty] | 2 |  |  | 0
Totals
[13 empty] | 78 |  |  | 0
Links to other AMR modules
- General AMR module
- Urine module
- Feces module
- Blood module
- Genital module
- Wound module
Yes
No
Yes
Partial
No
N/A
Yes
Partial
No
Yes
No
N/A

## Wound Module
AMR Lab Scorecard v9.2 (Sep 2021)
Bacterial culture, detection, identification and antimicrobial susceptibility testing of wound samples
W-A | How many wound culture tests and molecular tests were performed last year? | [12 empty] | <------- | It is highly recommended that assessors obtain the necessary permission to review the laboratory data. However, if assessors are unable to review the laboratory this question is NOT compulsory for completion of the assessment.
[5 empty] | Culture |  |  |  | Molecular
[5 empty] | Q1 | Q2 | Q3 | Q4 | Q1 | Q2 | Q3 | Q4 | Sub total | <------- | See: http://www.who.int/glass/en/ and other frequently isolated pathogens.
| Hospital-acquired | [12 empty] | <------- | Hospital-acquired infections are defined as bacterial infections in hospitalized patients (i.e. pathogenic bacterial isolated from a sample collected more than 48 hours after admission).
|  | S. aureus | [10 empty] | 0
|  | S. pyogenes | [10 empty] | 0
|  | Enterococcus sp. | [10 empty] | 0
|  | Enterobacteriaceae | [10 empty] | 0
|  | P. aeruginosa | [10 empty] | 0
| Community-acquired | [12 empty] | <------- | Community-acquired infections are defined as ambulatory patients and hospitalized patients from which a sample was collected less than 48 hours after admission.
|  | S. aureus | [10 empty] | 0
|  | S. pyogenes | [10 empty] | 0
|  | Enterococcus sp. | [10 empty] | 0
|  | Enterobacteriaceae | [10 empty] | 0
|  | P. aeruginosa | [10 empty] | 0
| Unknown/referred | [12 empty] | <------- | If the laboratory can’t distinguish between hospital & community acquired infections, the number of organisms isolated should be recorded as "Unknown/referred".
|  | S. aureus | [10 empty] | 0
|  | S. pyogenes | [10 empty] | 0
|  | Enterococcus sp. | [10 empty] | 0 |  | Total number of isolates per pathogen
|  | Enterobacteriaceae | [10 empty] | 0
|  | P. aeruginosa | [10 empty] | 0 |  | S. aureus |  | 0
| TOTAL ISOLATES |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | S. pyogenes |  | 0
| TOTAL NUMBER OF WOUND TESTS PERFORMED | [13 empty] | Enterococcus sp. |  | 0
| TOTAL NUMBER OF WOUND CULTURES WITH NO PATHOGENS ISOLATED OR IDENTIFIED | [13 empty] | Enterobacteriaceae |  | 0
[15 empty] | P. aeruginosa |  | 0
Q = Quarter
W-B | Are there any significant variations (> 20%) in the number of wound cultures or other tests performed or organisms isolated or identified each quarter? If ‘Yes’, please explain.
[13 empty] | <------- | It is highly recommended that assessors obtain the necessary permission to review the laboratory data. However, if assessors are unable to review the laboratory this question is NOT compulsory for completion of the assessment
Section 1: Documents and records
All generic requirements apply, see SLIPTA Section 1. In addition, assessors should review the following:
Question no. |  | Question | [10 empty] | Available score |  | Answer | Score | Comments | Related to SLIPTA question
W1.1 |  | Does the laboratory have documentation covering the following processes? | [10 empty] | 2 |  | Answer all sub questions | 0 |  | 1.5
|  |  | a) Production of Blood Agar, MacConkey Agar or other media for wound culture pathogen isolation? | [15 empty] | 100
|  |  | b) Processing of wound samples | [15 empty] | 100
|  |  | c) Detection, identification and AST of wound pathogens | [15 empty] | 100
|  |  | d) Reporting of wound culture and molecular test results | [15 empty] | 100
|  |  | e) Interlaboratory comparison or proficiency testing (PT) | [15 empty] | 100
|  |  | f) Laboratory safety | [15 empty] | 100 | 600
W1.2 |  | Are the documents complete, in-date and witnessed by all staff performing wound culture and molecular tests? | [10 empty] | 2 |  |  | 0 |  | 1.5
See ISO15189:2012 Clause 5.5.3 for minimum requirements for a technical Standard Operating Procedure (SOP).
W1.3 |  | Are/is the following processes documented? | [10 empty] | 3 |  | Answer all sub questions | 0 |  | 1.5
|  |  | a) Rejection criteria for wound samples? | [15 empty] | 100
|  |  | b) A policy for reporting critical results? | [15 empty] | 100
|  |  | c) Instructions for reporting wound culture tests with mixed bacterial growth? | [15 empty] | 100
|  |  | d) Instructions for referral of wound culture or molecular tests at the laboratory? | [15 empty] | 100
|  |  | e) Instructions for handling samples received after hours? | [15 empty] | 100
|  |  | f) Instructions for referral of bacterial isolates for identification and AST? | [15 empty] | 100
|  |  | g) Instructions on how to perform AST conversions for automated, disk diffusion, Etest / Gradient and microdilution AST? | [15 empty] | 100
|  |  | h) Turnaround time for wound culture or molecular tests ? | [15 empty] | 100
|  |  | i) Definition of rare / unexpected AST results? | [15 empty] | 100
|  |  | j) Confirmatory tests for unusual or unexpected patient AST results? | [15 empty] | 100 | 1000
Sub total | [12 empty] | 7 |  |  | 0
Section 2: Management Reviews
All generic requirements apply, see SLIPTA Section 2.
Section 3: Organization & Personnel
All generic requirements apply, see SLIPTA Section 3. In addition, assessors should review the following:
Question no. |  | Question | [10 empty] | Available score |  | Answer | Score | Comments | Related to SLIPTA question
W3.1 |  | Is there evidence that laboratory staff have been trained and are competent in the following: | [10 empty] | 3 |  | Answer all sub questions | 0 |  | 3.6
|  |  | a) Processing of wound samples for culture and molecular tests | [15 empty] | 100
|  |  | b) Identification and AST of wound pathogens | [15 empty] | 100
|  |  | c) Interpretation of wound culture and molecular test results | [15 empty] | 100
|  |  | d) Reporting of wound culture and molecular test results | [15 empty] | 100
|  |  | e) QC, EQA & PT for wound culture and molecular tests | [15 empty] | 100
|  |  | f) Laboratory safety | [15 empty] | 100 | 600
Review training records, competency assessment forms and duty rosters. Pay attention to date of training and scope of training compared with techniques being performed.
W3.2 |  | Is there evidence that laboratory staff are following the procedures described in the laboratory documentation? | [10 empty] | 3 |  | Answer all sub questions | 0 |  | 3.7
|  |  | a) Processing of wound for culture and molecular tests | [15 empty] | 100
|  |  | b) Interpretation of wound culture test results | [15 empty] | 100
|  |  | c) Identification and AST of wound pathogens | [15 empty] | 100
|  |  | d) Reporting of wound culture test and molecular test results | [15 empty] | 100 | 400
Directly observe procedures being performed compared to the SOP.
Sub total | [12 empty] | 6 |  |  | 0
Section 4: Client Management & Customer Service
All generic requirements apply, see SLIPTA Section 4. In addition, assessors should review the following:
Question no. |  | Question | [10 empty] | Available score |  | Answer | Score | Comments | Related to SLIPTA question
W4.1 |  | Is there evidence that the laboratory has provided clients information / instructions on wound sample collection, storage and transportation to the laboratory? Does the information / instructions include: | [10 empty] | 3 |  | Answer all sub questions | 0 |  | 4.1
|  |  | a) Use of sterile techniques for collecting wound samples from sterile sites? | [15 empty] | 100
|  |  | b) Collection procedures for culture of anaerobic organisms? | [15 empty] | 100
|  |  | c) Interpretation of contaminated results? | [15 empty] | 100 | 300
W4.2 |  | Is there evidence that the laboratory has provided clients information / instructions on interpretation of wound culture results and AST? | [10 empty] | 2 |  |  | 0 |  | 4.1
Sub total | [12 empty] | 5 |  |  | 0
Section 5: Equipment
All generic requirements apply, see SLIPTA Section 5.
Section 6: Evaluation and Audits
All generic requirements apply, see SLIPTA Section 6. Note: this checklist (with or without SLIPTA) can be used for internal and external audits.
Section 7: Purchasing & Inventory
All generic requirements apply, see SLIPTA Section 7. In addition, assessors should review the following:
Question no. |  | Question | [10 empty] | Available score |  | Answer | Score | Comments | Related to SLIPTA question
W7.1 |  | Is all media for bacterial culture isolation, identification and AST stored correctly and in date (from date of manufacture media must be stored at 2-8 °C)? According to manufacturer's requirements. | [10 empty] | 2 |  | Answer all sub questions | 0 |  | 7.10
|  |  | - Blood agar | [15 empty] | 100
|  |  | - MacConkey agar | [15 empty] | 100
|  |  | - Mueller Hinton | [15 empty] | 100
|  |  | - CNA agar | [15 empty] | 100 | 400
Sub total | [12 empty] | 2 |  |  | 0
Section 8: Process Control
All generic requirements apply, see SLIPTA Section 8. In addition, assessors should review the following:
Question no. |  | Question | [10 empty] | Available score |  | Answer | Score | Comments | Related to SLIPTA question
MEDIA QUALITY CONTROL
W8.1 |  | Does the laboratory perform QC testing on all media before use? This includes in-house made or purchased from commercial sources. | [10 empty] | 3 |  | Answer all sub questions | 0 |  | 8.8
Blood agar
|  |  | Do QC records for blood agar plates demonstrate that they are checked for their ability to support growth of fastidious organisms such as S. pneumoniae? | [15 empty] | 100
|  |  | Do QC records for blood agar plates demonstrate that they are checked for their ability to show beta, alpha, and gamma hemolysis? | [15 empty] | 100
MacConkey agar (MAC)
|  |  | Do QC records for MAC plates demonstrate that they are checked for their ability to suppress growth of Gram-positive organisms while allowing the growth of Gram-negative organisms? | [15 empty] | 100
|  |  | Do QC records for MAC plates demonstrate that they are checked for their ability to allow visualization of lactose fermentation? | [15 empty] | 100
Mueller Hinton Ager (MHA)
|  |  | Do QC records demonstrate that MHA plates are checked for their ability to grow S. aureus & E. coli? | [15 empty] | 100
CNA agar
|  |  | Do QC records for CNA agar plates demonstrate that they are checked for their ability to suppress growth of Gram-negative organisms while allowing the growth of Gram-positive organisms? | [15 empty] | 100
|  |  | Do QC records demonstrate that CNA agar plates are checked for their ability to grow S. pyogenes & S. aureus? | [15 empty] | 100 | 700
W8.2 |  | Does the laboratory: | [10 empty] | 2 |  | Answer all sub questions | 0 |  | 8.8
|  |  | Perform sterility and performance tests for every batch of culture media using certified reference strains as controls? | [15 empty] | 100
|  |  | Are reference strains sourced from an authorized supplier (e.g. ATCC)? | [15 empty] | 100
|  |  | Are the reference strains stored, cultured and sub-cultured in accordance with the specification from the supplier? | [15 empty] | 100 | 300
W8.3 |  | Does the laboratory determine the cause of failed QC (root cause analysis), perform corrective actions and measure the effectiveness thereof? | [10 empty] | 3 |  |  | 0 |  | 8.10 | 100
WOUND CULTURE PROCEDURE
For complete recommended procedure, see the User Guide.
W8.4 |  | Are wound cultures plated onto non-selective media including (at least): | [10 empty] | 2 |  | Answer all sub questions | 0 |  | 8.7
|  |  | Blood agar | [15 empty] | 100
|  |  | MacConkey agar | [15 empty] | 100 | 200
W8.5 |  | Are wound culture plates incubated at 35-37 degrees Celsius aerobically and anaerobically if applicable? | [10 empty] | 2 |  |  | 0 |  | 8.7
W8.6 |  | Does the laboratory report wound cultures as contaminated if they contain organisms that should be considered contaminants (e.g. Bacillus sp., Coagulase-negative Staphylococcus, Corynebacterium sp.)? | [10 empty] | 2 |  |  | 0 |  | 8.7
BACTERIAL ID & AST
W8.7 |  | Is the following testing performed for S. aureus identification: If the laboratory performs penicillin AST, it is recommended that S. aureus isolates with penicillin zones sizes or MICs in the susceptible range are tested for B-lactamase production using the zone-edge test or a nitrocefin test before being reported as penicillin susceptible. | [10 empty] | 2 |  | Answer all sub questions | 0 |  | 8.7
|  |  | Catalase | [15 empty] | 100
|  |  | Coagulase (slide or tube) | [15 empty] | 100
|  |  | Mannitol Salt Agar (MSA) | [15 empty] | 100
|  |  | Dnase | [15 empty] | 100 | 400
W8.8 |  | Does S. aureus AST include the following antibiotics: If oxacillin and cefoxitin results are discrepant for S. aureus (one is susceptible and one is resistant), the laboratory should repeat the testing. Note: oxacillin testing should always be tested by MIC (not disc diffusion). If the results remain discrepant, oxacillin should be reported as resistant. | [10 empty] | 2 |  | Answer all sub questions | 0 |  | 8.7
|  |  | Cefoxitin | [15 empty] | 100
|  |  | Vancomycin | [15 empty] | 100 | 200
W8.9 |  | Does the laboratory detect methicillin/nafcillin resistance in S. aureus using oxacillin disk? | [10 empty] | 2 |  |  | 0 |  | 8.7 | 100
W8.10 |  | Is the following testing performed for Streptococcus sp. and Enterococcus sp. identification: | [10 empty] | 2 |  | Answer all sub questions | 0 |  | 8.7
|  |  | Bacitracin | [15 empty] | 100
|  |  | Pyrrolidonyl Arylamidase (PYR) | [15 empty] | 100
|  |  | Bile solubility | [15 empty] | 100
|  |  | Optochin | [15 empty] | 100
|  |  | S. pneumoniae latex | [15 empty] | 100 | 400
W8.11 |  | Does Streptococcus sp. AST include the following antibiotics: | [10 empty] | 2 |  | Answer all sub questions | 0 |  | 8.7
|  |  | - Oxacillin If the laboratory uses an oxacillin disk (1ug) to screen for penicillin resistance (Penicillin G or Benzylpenicillin, the IV formulation) in S. pneumoniae and the zone size < 20, then the laboratory must do an MIC method before reporting penicillin as resistant (CLSI recommendation). EUCAST recommends that if the zone size is < 20mm to do a MIC, if ≧ 20 mm the result should be reported as susceptible. | [15 empty] | 100
|  |  | Co-trimoxazole | [15 empty] | 100
|  |  | Ceftriaxone or cefotaxime | [15 empty] | 100 | 300
W8.12 |  | Is the following testing performed to identify Gram negative bacilli: | [10 empty] | 2 |  | Answer all sub questions | 0 |  | 8.7
|  |  | Oxidase | [15 empty] | 100
|  |  | Indole | [15 empty] | 100
|  |  | Methyl Red | [15 empty] | 100
|  |  | Voges Proskauer | [15 empty] | 100
|  |  | Citrate | [15 empty] | 100
|  |  | Triple Sugar Iron or Kligler Iron | [15 empty] | 100
|  |  | Urease | [15 empty] | 100
|  |  | Motility | [15 empty] | 100 | 800
W8.13 |  | Does the lab follow the latest CLSI/EUCAST guidelines for AST of Gram negative bacilli? https://www.clsi.org/ www.eucast.org/ | [10 empty] | 2 |  |  | 0 |  | 8.7 | 100
W8.14 |  | Does the laboratory use Combination Disk Test or another equivalent method for Extended Spectrum Beta-Lactamase (ESBL) screening? J Clin Microbiol. 2013 Sep; 51(9): 2986–2990. | [10 empty] | 2 |  |  | 0 |  | 8.7 | 100
W8.15 |  | Does the laboratory use Combination Disk Test or another equivalent method for carbapenemase screening? | [10 empty] | 2 |  |  | 0 |  | 8.7 | 100
INTERLABORATORY COMPARISON, PT AND EXTERNAL QUALITY ASSURANCE (EQA)
W8.16 |  | Is the laboratory enrolled in an interlaboratory comparison or PT program for wound culture and molecular tests for organism identification, and AST? | [10 empty] | 2 |  |  | 0 |  | 8.14
W8.17 |  | Did the laboratory pass the last 3 rounds of interlaboratory comparison or PT program testing? | [10 empty] | 2 |  |  | 0 |  | 8.14
W8.18 |  | Does the laboratory receive onsite supervision visits as part of the EQA program for wound culture and molecular tests? | [10 empty] | 2 |  |  | 0 |  | 8.14
Sub total | [12 empty] | 38 |  |  | 0
Section 9: Information Management
All generic requirements apply, see SLIPTA Section 9. In addition, assessors should review the following:
Question no. |  | Question | [10 empty] | Available score |  | Answer | Score | Comments | Related to SLIPTA question
W9.1 |  | Does the final report for wound culture list the organisms for which the specimen was and was not cultured? The laboratory should inform the clinician on the report what organisms were excluded during the culture process. This may be either by choice of media or incubation conditions (e.g. anaerobic organisms). Assessors should review a number of laboratory reports to determine how results are reported. Procedures should be consistent with the laboratory’s SOPs. | [10 empty] | 2 |  |  | 0 |  | 9.3
W9.2 |  | Does the laboratory report alert organisms which include at least? Alert organisms are organisms with significant public health threat and / or organisms that are notifiable. | [10 empty] | 2 |  | Answer all sub questions | 0 |  | 9.3
|  |  | - Methicillin resistant S. aureus | [15 empty] | 100
|  |  | - Imipenem resistant Enterobacteriaceae | [15 empty] | 100
|  |  | - Carbapenem resistant Enterobacteriaceae | [15 empty] | 100
|  |  | - ESBL producing organisms | [15 empty] | 100
|  |  | - Multidrug resistant Pseudomonas | [15 empty] | 100 | 500
Sub total | [12 empty] | 4 |  |  | 0
Section 10: Identification of Nonconformities, Corrective and Preventive Actions
All generic requirements apply, see SLIPTA Section 10.
Section 11: Occurrence/Incident Management & Process Improvement
All generic requirements apply, see SLIPTA Section 11. In addition, assessors should review the following:
Question no. |  | Question | [10 empty] | Available score |  | Answer | Score | Comments | Related to SLIPTA question
W11.1 |  | Are the following performance indicators collected? It may not be possible for laboratories to distinguish between community and hospital acquired infection if this is not collected on the laboratory requisition form. | [10 empty] | 3 |  | Answer all sub questions | 0 |  | 11.4 & 11.5
|  |  | Number of wound culture and molecular tests performed (disaggregated by type): | [11 empty] | Answer all sub questions |  |  |  | 100
|  |  |  | Hospital-acquired Hospital-acquired infections are defined as bacterial infections in hospitalized patients (i.e. pathogenic bacterial isolated from a sample collected more than 48 hours after admission). | [14 empty] | 100
|  |  |  | Community-acquired Community-acquired infections are defined as ambulatory patients and hospitalized patients from which a sample was collected less than 48 hours after admission. | [14 empty] | 100
|  |  |  | Unknown/referred If the laboratory can’t distinguish between hospital & community acquired infections, the number of organisms isolated should be recorded as "Unknown/referred". | [14 empty] | 100 | 300
|  |  | Number of wound culture and molecular tests where pathogens were isolated (disaggregated by type): | [11 empty] | Answer all sub questions |  |  |  | 100
|  |  |  | S. aureus | [14 empty] | 100
|  |  |  | S. pyogenes | [14 empty] | 100
|  |  |  | Enterococcus sp. | [14 empty] | 100
|  |  |  | Enterobacteriaceae | [14 empty] | 100
|  |  |  | P. aeruginosa | [14 empty] | 100 | 500
|  |  | Total number of wound cultures with no pathogens Isolated or identified | [15 empty] | 100
|  |  | Wound culture and molecular test TAT (disaggregated by in-patient & out-patient and by type) | [15 empty] | 100 | 400
Sub total | [12 empty] | 3 |  |  | 0
Section 12: Facilities and Biosafety
All generic requirements apply, see SLIPTA Section 12.
Totals
[13 empty] | 65 |  |  | 0
Links to other AMR modules
- General AMR module
- Urine module
- Feces module
- Blood module
- Genital module
- Pulmonary module
Yes
No
Yes
Partial
No
N/A
Yes
Partial
No
Yes
No
N/A

## Previous Audit Information
AMR Lab Scorecard v9.2 (Sep 2021)
Previous audit information
Prior AMR audit status | Baseline Audit |  | Audit 1 |  | Audit 2 |  | Audit 3
Date of previous audit
Please provide values for each section | Score | Percentage | Score | Percentage | Score | Percentage | Score | Percentage
General scorecard
1. Documents & Records
2. Management Reviews
3. Organization & Personnel
4. Client Management & Customer Service
5. Equipment
6. Internal Audit
7. Purchasing & Inventory
8. Process Control, IQA and EQA
9. Information Management
10. Corrective Action
11. Occurrence Management & Process Improvement
12. Facilities & Safety
Overall general scorecard
Urine scorecard
1. Documents & Records
2. Management Reviews
3. Organization & Personnel
4. Client Management & Customer Service
5. Equipment
6. Internal Audit
7. Purchasing & Inventory
8. Process Control, IQA and EQA
9. Information Management
10. Corrective Action
11. Occurrence Management & Process Improvement
12. Facilities & Safety
Overall urine scorecard
Feces scorecard
1. Documents & Records
2. Management Reviews
3. Organization & Personnel
4. Client Management & Customer Service
5. Equipment
6. Internal Audit
7. Purchasing & Inventory
8. Process Control, IQA and EQA
9. Information Management
10. Corrective Action
11. Occurrence Management & Process Improvement
12. Facilities & Safety
Overall feces scorecard
Blood scorecard
1. Documents & Records
2. Management Reviews
3. Organization & Personnel
4. Client Management & Customer Service
5. Equipment
6. Internal Audit
7. Purchasing & Inventory
8. Process Control, IQA and EQA
9. Information Management
10. Corrective Action
11. Occurrence Management & Process Improvement
12. Facilities & Safety
Overall blood scorecard
Genital scorecard
1. Documents & Records
2. Management Reviews
3. Organization & Personnel
4. Client Management & Customer Service
5. Equipment
6. Internal Audit
7. Purchasing & Inventory
8. Process Control, IQA and EQA
9. Information Management
10. Corrective Action
11. Occurrence Management & Process Improvement
12. Facilities & Safety
Overall genital scorecard
Pumonary scorecard
1. Documents & Records
2. Management Reviews
3. Organization & Personnel
4. Client Management & Customer Service
5. Equipment
6. Internal Audit
7. Purchasing & Inventory
8. Process Control, IQA and EQA
9. Information Management
10. Corrective Action
11. Occurrence Management & Process Improvement
12. Facilities & Safety
Overall pulmonary scorecard
Wound scorecard
1. Documents & Records
2. Management Reviews
3. Organization & Personnel
4. Client Management & Customer Service
5. Equipment
6. Internal Audit
7. Purchasing & Inventory
8. Process Control, IQA and EQA
9. Information Management
10. Corrective Action
11. Occurrence Management & Process Improvement
12. Facilities & Safety
Overall wound scorecard
Prior SLIPTA audit status | Baseline Audit |  | Audit 1 |  | Audit 2 |  | Audit 3
Date of previous audit
Prior audit star rating
Please provide values for each section | Score | Percentage | Score | Percentage | Score | Percentage | Score | Percentage
1. Documents & Records
2. Management Reviews
3. Organization & Personnel
4. Client Management & Customer Service
5. Equipment
6. Internal Audit
7. Purchasing & Inventory
8. Process Control, IQA and EQA
9. Information Management
10. Corrective Action
11. Occurrence Management & Process Improvement
12. Facilities & Safety
Overall

## AMR SUMMARY REPORT
AMR Lab Scorecard v9.2 (Sep 2021)
AMR Summary Report
LABORATORY DETAILS
Laboratory name:
Laboratory address:
Laboratory telephone number:
Laboratory email:
Laboratory level:
Laboratory type:
Contact person of laboratory:
Telephone number:
AUDIT DETAILS
Audit carried out by:
Audit date:
AMR AUDIT SCORE SUMMARY
To compare with previous audit data, select previous audit: |  |  |  | Select previous audit
General scorecard
|  | Percentage score |  |  | Number of stars
Current audit: |  | 0% |  |  | 0 stars
Previous audit
|  | Urine scorecard |  |  | Feces scorecard
|  | Percentage score | Number of stars |  | Percentage score |  | Number of stars
Current audit:
Previous audit
|  | Blood scorecard |  |  | Genital scorecard
|  | Percentage score | Number of stars |  | Percentage score |  | Number of stars
Current audit:
Previous audit
|  | Pulmonary scorecard |  |  | Wound scorecard
|  | Percentage score | Number of stars |  | Percentage score |  | Number of stars
Current audit:
Previous audit
Average
|  | Percentage score |  |  | Number of stars
Current audit: |  | 0% |  |  | 0 stars
Previous audit
GENERAL MODULE AUDIT SCORE SUMMARY
Section |  |  |  |  | Total points possible | Current Audit |  | Previous audit | Change in score since previous audit
[6 empty] | Date: |  | Date:
[6 empty] | Current audit score |  | Previous audit score
| 1. Documents and Records |  |  |  | 11 | 0 | 0%
| 2. Management Reviews |  |  |  | 8 | 0 | 0%
| 3. Organization and Personnel |  |  |  | N/A | N/A | N/A
| 4. Client Management and Customer Service |  |  |  | 2 | 0 | 0%
| 5. Equipment |  |  |  | 10 | 0 | 0%
| 6. Evaluation and Audits |  |  |  | 5 | 0 | 0%
| 7. Purchasing and Inventory |  |  |  | 4 | 0 | 0%
| 8. Process Control and Internal and External Quality Assessment |  |  |  | 17 | 0 | 0%
| 9. Information Management |  |  |  | 2 | 0 | 0%
| 10. Corrective Action |  |  |  | 8 | 0 | 0%
| 11. Occurrence Management and Process Improvement |  |  |  | 4 | 0 | 0%
| 12. Facilities and Safety |  |  |  | 4 | 0 | 0%
General module total |  |  |  |  | 75 | 0 | 0%
N/A | 1. Documents & Records |  | S1
Baseline audit | 2. Management Reviews |  | S2
Audit 1 | 3. Organization & Personnel |  | S3
Audit 2 | 4. Client Management & Customer Service |  | S4
Audit 3 | 5. Equipment |  | S5
| 6. Evaluation & Audits |  | S6
| 7. Purchasing & Inventory |  | S7
| 8. Process Control, IQA & EQA |  | S8
| 9. Information Management |  | S9
| 10. Corrective Action |  | S10
| 11. Occurrence Management & Process Improvement |  | S11
| 12. Facilities & Safety |  | S12
URINE MODULE AUDIT SCORE SUMMARY
Section |  |  |  |  | Total points possible | Current Audit |  | Previous audit | Change in score since previous audit
[6 empty] | Date: |  | Date:
[6 empty] | Current audit score |  | Previous audit score
| 1. Documents and Records |  |  |  | 7 | 0 | 0%
| 2. Management Reviews |  |  |  | N/A | N/A | N/A
| 3. Organization and Personnel |  |  |  | 6 | 0 | 0%
| 4. Client Management and Customer Service |  |  |  | 5 | 0 | 0%
| 5. Equipment |  |  |  | N/A | N/A | N/A
| 6. Evaluation and Audits |  |  |  | N/A | N/A | N/A
| 7. Purchasing and Inventory |  |  |  | 2 | 0 | 0%
| 8. Process Control and Internal and External Quality Assessment |  |  |  | 52 | 0 | 0%
| 9. Information Management |  |  |  | 4 | 0 | 0%
| 10. Corrective Action |  |  |  | N/A | N/A | N/A
| 11. Occurrence Management and Process Improvement |  |  |  | 3 | 0 | 0%
| 12. Facilities and Safety |  |  |  | N/A | N/A | N/A
Urine module total |  |  |  |  | 79 | 0 | 0%
N/A | 1. Documents & Records |  | S1
Baseline audit | 2. Management Reviews |  | S2
Audit 1 | 3. Organization & Personnel |  | S3
Audit 2 | 4. Client Management & Customer Service |  | S4
Audit 3 | 5. Equipment |  | S5
| 6. Evaluation & Audits |  | S6
| 7. Purchasing & Inventory |  | S7
| 8. Process Control, IQA & EQA |  | S8
| 9. Information Management |  | S9
| 10. Corrective Action |  | S10
| 11. Occurrence Management & Process Improvement |  | S11
| 12. Facilities & Safety |  | S12
FECES MODULE AUDIT SCORE SUMMARY
Section |  |  |  |  | Total points possible | Current Audit |  | Previous audit | Change in score since previous audit
[6 empty] | Date: |  | Date:
[6 empty] | Current audit score |  | Previous audit score
| 1. Documents and Records |  |  |  | 7 | 0 | 0%
| 2. Management Reviews |  |  |  | N/A | N/A | N/A
| 3. Organization and Personnel |  |  |  | 6 | 0 | 0%
| 4. Client Management and Customer Service |  |  |  | 5 | 0 | 0%
| 5. Equipment |  |  |  | N/A | N/A | N/A
| 6. Evaluation and Audits |  |  |  | N/A | N/A | N/A
| 7. Purchasing and Inventory |  |  |  | 2 | 0 | 0%
| 8. Process Control and Internal and External Quality Assessment |  |  |  | 44 | 0 | 0%
| 9. Information Management |  |  |  | 4 | 0 | 0%
| 10. Corrective Action |  |  |  | N/A | N/A | N/A
| 11. Occurrence Management and Process Improvement |  |  |  | 3 | 0 | 0%
| 12. Facilities and Safety |  |  |  | N/A | N/A | N/A
Feces module total |  |  |  |  | 71 | 0 | 0%
Baseline audit | 2. Management Reviews |  | S2
Audit 1 | 3. Organization & Personnel |  | S3
Audit 2 | 4. Client Management & Customer Service |  | S4
Audit 3 | 5. Equipment |  | S5
| 6. Evaluation & Audits |  | S6
| 7. Purchasing & Inventory |  | S7
| 8. Process Control, IQA & EQA |  | S8
| 9. Information Management |  | S9
| 10. Corrective Action |  | S10
| 11. Occurrence Management & Process Improvement |  | S11
| 12. Facilities & Safety |  | S12
BLOOD MODULE AUDIT SCORE SUMMARY
Section |  |  |  |  | Total points possible | Current Audit |  | Previous audit | Change in score since previous audit
[6 empty] | Date: |  | Date:
[6 empty] | Current audit score |  | Previous audit score
| 1. Documents and Records |  |  |  | 7 | 0 | 0%
| 2. Management Reviews |  |  |  | N/A | N/A | N/A
| 3. Organization and Personnel |  |  |  | 6 | 0 | 0%
| 4. Client Management and Customer Service |  |  |  | 5 | 0 | 0%
| 5. Equipment |  |  |  | N/A | N/A | N/A
| 6. Evaluation and Audits |  |  |  | N/A | N/A | N/A
| 7. Purchasing and Inventory |  |  |  | 2 | 0 | 0%
| 8. Process Control and Internal and External Quality Assessment |  |  |  | 45 | 0 | 0%
| 9. Information Management |  |  |  | 4 | 0 | 0%
| 10. Corrective Action |  |  |  | N/A | N/A | N/A
| 11. Occurrence Management and Process Improvement |  |  |  | 3 | 0 | 0%
| 12. Facilities and Safety |  |  |  | 4 | 0 | 0%
Blood module total |  |  |  |  | 76 | 0 | 0%
N/A | 1. Documents & Records |  | S1
Baseline audit | 2. Management Reviews |  | S2
Audit 1 | 3. Organization & Personnel |  | S3
Audit 2 | 4. Client Management & Customer Service |  | S4
Audit 3 | 5. Equipment |  | S5
| 6. Evaluation & Audits |  | S6
| 7. Purchasing & Inventory |  | S7
| 8. Process Control, IQA & EQA |  | S8
| 9. Information Management |  | S9
| 10. Corrective Action |  | S10
| 11. Occurrence Management & Process Improvement |  | S11
| 12. Facilities & Safety |  | S12
GENITAL MODULE AUDIT SCORE SUMMARY
Section |  |  |  |  | Total points possible | Current Audit |  | Previous audit | Change in score since previous audit
[6 empty] | Date: |  | Date:
[6 empty] | Current audit score |  | Previous audit score
| 1. Documents and Records |  |  |  | 7 | 0 | 0%
| 2. Management Reviews |  |  |  | N/A | N/A | N/A
| 3. Organization and Personnel |  |  |  | 6 | 0 | 0%
| 4. Client Management and Customer Service |  |  |  | 5 | 0 | 0%
| 5. Equipment |  |  |  | N/A | N/A | N/A
| 6. Evaluation and Audits |  |  |  | N/A | N/A | N/A
| 7. Purchasing and Inventory |  |  |  | 2 | 0 | 0%
| 8. Process Control and Internal and External Quality Assessment |  |  |  | 39 | 0 | 0%
| 9. Information Management |  |  |  | 2 | 0 | 0%
| 10. Corrective Action |  |  |  | N/A | N/A | N/A
| 11. Occurrence Management and Process Improvement |  |  |  | 3 | 0 | 0%
| 12. Facilities and Safety |  |  |  | N/A | N/A | N/A
Genital module total |  |  |  |  | 64 | 0 | 0%
N/A | 1. Documents & Records |  | S1
Baseline audit | 2. Management Reviews |  | S2
Audit 1 | 3. Organization & Personnel |  | S3
Audit 2 | 4. Client Management & Customer Service |  | S4
Audit 3 | 5. Equipment |  | S5
| 6. Evaluation & Audits |  | S6
| 7. Purchasing & Inventory |  | S7
| 8. Process Control, IQA & EQA |  | S8
| 9. Information Management |  | S9
| 10. Corrective Action |  | S10
| 11. Occurrence Management & Process Improvement |  | S11
| 12. Facilities & Safety |  | S12
PULMONARY MODULE AUDIT SCORE SUMMARY
Section |  |  |  |  | Total points possible | Current Audit |  | Previous audit | Change in score since previous audit
[6 empty] | Date: |  | Date:
[6 empty] | Current audit score |  | Previous audit score
| 1. Documents and Records |  |  |  | 7 | 0 | 0%
| 2. Management Reviews |  |  |  | N/A | N/A | N/A
| 3. Organization and Personnel |  |  |  | 6 | 0 | 0%
| 4. Client Management and Customer Service |  |  |  | 5 | 0 | 0%
| 5. Equipment |  |  |  | N/A | N/A | N/A
| 6. Evaluation and Audits |  |  |  | N/A | N/A | N/A
| 7. Purchasing and Inventory |  |  |  | 2 | 0 | 0%
| 8. Process Control and Internal and External Quality Assessment |  |  |  | 49 | 0 | 0%
| 9. Information Management |  |  |  | 4 | 0 | 0%
| 10. Corrective Action |  |  |  | N/A | N/A | N/A
| 11. Occurrence Management and Process Improvement |  |  |  | 3 | 0 | 0%
| 12. Facilities and Safety |  |  |  | 2 | 0 | 0%
Pulmonary module total |  |  |  |  | 78 | 0 | 0%
N/A | 1. Documents & Records |  | S1
Baseline audit | 2. Management Reviews |  | S2
Audit 1 | 3. Organization & Personnel |  | S3
Audit 2 | 4. Client Management & Customer Service |  | S4
Audit 3 | 5. Equipment |  | S5
| 6. Evaluation & Audits |  | S6
| 7. Purchasing & Inventory |  | S7
| 8. Process Control, IQA & EQA |  | S8
| 9. Information Management |  | S9
| 10. Corrective Action |  | S10
| 11. Occurrence Management & Process Improvement |  | S11
| 12. Facilities & Safety |  | S12
WOUND MODULE AUDIT SCORE SUMMARY
Section |  |  |  |  | Total points possible | Current Audit |  | Previous audit | Change in score since previous audit
[6 empty] | Date: |  | Date:
[6 empty] | Current audit score |  | Previous audit score
| 1. Documents and Records |  |  |  | 7 | 0 | 0%
| 2. Management Reviews |  |  |  | N/A | N/A | N/A
| 3. Organization and Personnel |  |  |  | 6 | 0 | 0%
| 4. Client Management and Customer Service |  |  |  | 5 | 0 | 0%
| 5. Equipment |  |  |  | N/A | N/A | N/A
| 6. Evaluation and Audits |  |  |  | N/A | N/A | N/A
| 7. Purchasing and Inventory |  |  |  | 2 | 0 | 0%
| 8. Process Control and Internal and External Quality Assessment |  |  |  | 38 | 0 | 0%
| 9. Information Management |  |  |  | 4 | 0 | 0%
| 10. Corrective Action |  |  |  | N/A | N/A | N/A
| 11. Occurrence Management and Process Improvement |  |  |  | 3 | 0 | 0%
| 12. Facilities and Safety |  |  |  | N/A | N/A | N/A
Wound module total |  |  |  |  | 65 | 0 | 0%
N/A | 1. Documents & Records |  | S1
Baseline audit | 2. Management Reviews |  | S2
Audit 1 | 3. Organization & Personnel |  | S3
Audit 2 | 4. Client Management & Customer Service |  | S4
Audit 3 | 5. Equipment |  | S5
| 6. Evaluation & Audits |  | S6
| 7. Purchasing & Inventory |  | S7
| 8. Process Control, IQA & EQA |  | S8
| 9. Information Management |  | S9
| 10. Corrective Action |  | S10
| 11. Occurrence Management & Process Improvement |  | S11
| 12. Facilities & Safety |  | S12

## SLIPTA SUMMARY REPORT
AMR Lab Scorecard v9.2 (Sep 2021)
SLIPTA Summary Report
LABORATORY DETAILS
Laboratory name:
Laboratory number:
Laboratory address:
Laboratory telephone number:
Laboratory fax number:
Laboratory email:
Laboratory level:
Laboratory type:
Head of laboratory:
Telephone number:
AUDIT DETAILS
Audit carried out by:
Audit date:
OVERALL AUDIT SCORE
| Percentage score | Star rating
Current audit: | 0% | 0 star
Select previous audit |  |  | <-- Select the dropdown if previous audit data was entered
DETAILED AUDIT SCORE
Section |  |  | Total points possible | Current Audit |  | Previous audit | Change in score since previous audit
|  |  |  | Audit date: |  | Audit date:
|  |  |  | Current audit score |  | Previous audit score
1. Documents and Records |  |  | 28 | 0 | 0%
2. Management Reviews |  |  | 14 | 0 | 0%
3. Organization and Personnel |  |  | 22 | 0 | 0%
4. Client Management and Customer Service |  |  | 10 | 0 | 0%
5. Equipment |  |  | 35 | 0 | 0%
6. Evaluation and Audits |  |  | 15 | 0 | 0%
7. Purchasing and Inventory |  |  | 24 | 0 | 0%
8. Process Control and Internal and External Quality Assessment |  |  | 32 | 0 | 0%
9. Information Management |  |  | 21 | 0 | 0%
10. Corrective Action |  |  | 19 | 0 | 0%
11. Occurrence Management and Process Improvement |  |  | 12 | 0 | 0%
12. Facilities and Safety |  |  | 43 | 0 | 0%
Total |  |  | 275 | 0 | 0%
N/A | 1. Documents & Records |  |  | S1
Baseline audit | 2. Management Reviews |  |  | S2
Audit 1 | 3. Organization & Personnel |  |  | S3
Audit 2 | 4. Client Management & Customer Service |  |  | S4
Audit 3 | 5. Equipment |  |  | S5
| 6. Evaluation & Audits |  |  | S6
| 7. Purchasing & Inventory |  |  | S7
| 8. Process Control, IQA & EQA |  |  | S8
| 9. Information Management |  |  | S9
| 10. Corrective Action |  |  | S10
| 11. Occurrence Management & Process Improvement |  |  | S11
| 12. Facilities & Safety |  |  | S12
| 13. AMR Technical proficiency |  |  | S13

## SLIPTA Laboratory Profile
SLIPTA Laboratory Profile
Site Information
Laboratory name:
Laboratory number:
Laboratory address:
Laboratory telephone number:
Laboratory fax number:
Laboratory email:
Laboratory level:
Laboratory type:
If other, please specify:
Head of Laboratory:
Telephone (Head of Laboratory):
Staff details
Number of degree holding staff: |  | Is this adequate?
Number of diploma holding staff: |  | Is this adequate?
Number of certificate holding staff: |  | Is this adequate?
Number of data clerks: |  | Is this adequate?
Number of phlebotomists: |  | Is this adequate?
Number of cleaners |  | Is this adequate?
Is the cleaner(s) dedicate to the laboratory only?
Has the cleaner(s) been trained in safe waste handling?
Number of drivers |  | Is this adequate?
Is the driver(s) dedicate to the laboratory only?
Has the driver(s) been trained in biosafety?
Number of other staff (specify positions below) |  | Is this adequate?
Supplemental details
Does the laboratory have sufficient space, equipment, supplies, personnel, infrastructure, etc. to execute the correct and timely performance of each test and maintain the quality management system?
Sufficient Space
Equipment
Supplies
Personnel
Infrastucture
Other (enter comments below)
National
Reference
Provincial
District
Zonal
Field
Public
Hospital
Private
Research
Non-hospital Outpatient Clinic
Other (please specify)
Yes
No
Insufficient Data
N/A
Baseline Audit
Audit 1
Audit 2
Audit 3
N/A
N/A
0 stars
1 star
2 stars
3 stars
4 stars
5 stars

## SLIPTA_S1
Section 1 - Documents and Records
Question no. | Question |  | Available Score | Answer | Score | Comments | Standard
1.1 | Legal Entity Does the laboratory have documentation stating its legal identity? |  | 2 |  | 0 |  | ISO15189:2012 Clause 4.1.1.2 The laboratory or the organization of which the laboratory is a part shall be an entity that can be held legally responsible for its activities. Note: Documentation could be in the form of a National Act, Company registration certificate, License number or Practice number.
1.2 | Laboratory Quality Manual Is there a current laboratory quality manual composed of the quality management system’s policies and procedures; has the manual content been communicated to staff, and is the manual understood and implemented by all staff? |  | 5 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 4.1.2.3 and 4.2.2.2 and 4.3 Note: A quality manual must be available that summarizes the laboratory’s quality management system, which includes policies that address all areas of the laboratory service, and identifies the goals and objectives of the quality management system. The quality manual must include policies and make reference to processes and procedures for all areas of the laboratory service and must address all the clauses of ISO15189:2012.
Does the quality manual includes the following elements:
|  | a) Quality policy statement that includes scope of service, standard of service, objectives of the quality management system, and management commitment to compliance | [5 empty] | 100
|  | b) Documented policies for the quality management system that meet the requirements of ISO15189:2012 (Refer to Question 1.5 of this checklist for list of policies required) | [5 empty] | 100
|  | c) Description of the quality management system and the structure of its documentation | [5 empty] | 100
|  | d) Reference to supporting procedures (SOPs), including managerial and technical procedures | [5 empty] | 100
|  | e) Description of the roles and responsibilities of the laboratory director, or laboratory manager, quality manager, and other key personnel (laboratory to define its key personnel) responsible for ensuring compliance | [5 empty] | 100
|  | f) Records of review and approval of the quality manual by authorized personnel | [5 empty] | 100
|  | g) Records to show that the quality manual was communicated to and understood by the lab personnel | [5 empty] | 100 | 700
1.3 | Document and Information Control System Does the laboratory have a system in place to control all documents and information (from internal and external sources)? |  | 2 |  | 0 |  | ISO15189:2012 Clause 4.3 Note: There must be a procedure on document control. A document control system must be in place to ensure that records and all documents (internal and external) are current, read and understood by personnel, approved by authorized persons, reviewed periodically and revised as required. Documents must be uniquely identified to include title, page numbers, and authority of issue, document number, versions, effective date, and author. Example of external documents includes regulations, standards, guidelines, equipment user manuals, package inserts, text books.
1.4 | Document and Records Is there a list that details all documents used in the quality management system indicating their editions and distribution? |  | 2 |  | 0 |  | ISO15189:2012 Clause 4.3 Note: Documents to be included on the list include Manuals, Procedures, and Processes. Work instructions. Forms, external documents.. The list could be in the form of a document master index, document log or document register. “Edition” can be regarded as synonymous with “revision or version” number for the documents.
1.5 | Laboratory Policies and Standard Operating Procedures Are policies and/or standard operating procedures (SOPs) for laboratory functions, technical and managerial procedures current, available and approved by authorized personnel? |  | 5 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 4.3 and 5.5 Note: The laboratory must define who is authorized to approve documents for its intended use. The approver should not be the author but can be the reviewer.
Has the laboratory defined Policies and/or SOPs that address the following:
|  | Ethical conduct How the laboratory will: 1) minimize activities that would diminish confidence in the laboratory's competence, impartiality, and judgment; 2) perform work within relevant legal requirements; 3) ensure confidentiality; 4) handle human samples, tissues or their remains as per regulations; 5) identify and avoid potential conflicts of interest and commercial, financial, political or other pressures that may affect the quality and integrity of operations? |  |  |  |  | ISO15189:2012 Clause 4.1.1.3 Note: Laboratories shall uphold the principle that the welfare and interest of the patient are paramount and patients should be treated fairly and without discrimination | 100
|  | Document Control How the laboratory will: 1) control all internal and external documents; 2) create documents; 3) identify documents; 4) review documents; 5) approve documents; 6) capture current versions and their distribution by means of a list; 7) handle amendments; 8) identify changes; 9) handle obsolete documents; 10) retain documents; 11) prevent the unintended use of any obsolete document; 12) ensure safe disposal of documents? |  |  |  |  | ISO15189:2012 Clause 4.3 and 4.13 Note: Documents that should be considered for document control are those that may vary based on changes in versions or time. Examples include policy statements, instructions for use, flow charts, procedures, specifications, forms, calibration tables, biological reference intervals and their origins, charts, posters, notices, memoranda, software documentation, drawings, plans, agreements, and documents of external origin such as regulations, standards and text books from which examination procedures are taken. | 100
|  | Control of Records How the laboratory will: 1) identify; 2) collect; 3) index; 4) access; 5) store; 6) maintain; 7) amends; 8) dispose of safely; 9) define the retention period for the identified records? |  |  |  |  | ISO15189 :2012 Clause 4.13 Note: Records can be in any form or type of medium providing they are readily accessible and protected from unauthorized alterations. Legal liability concerns regarding certain types of procedures (e.g. histology examinations, genetic examinations, pediatric examinations) may require the retention of certain records for much longer periods than for other records. For some records, especially those stored electronically, the safest storage may be on secure media and an offsite location. Type of records will include but not be limited to quality records, technical records, personnel records, test request and results records, | 100
|  | Communication (internal and external) How the laboratory will: 1) ensure effective communication with staff and users of the laboratory; 2) handle staff suggestions for improvement; 3) communicate with stakeholders on the effectiveness of the quality management system across all processes; 4) capture records of all communications; 5) retain and maintain all records of communication, requests, inquiries, verbal discussions and requests for additional examinations, meeting agendas, and meeting minutes)? |  |  |  |  | ISO15189:2012 Clause 4.1.2.6 and 4.14 Note: Laboratory management must ensure that appropriate communication processes are established between the laboratory and its stakeholders and that communication takes place regarding the effectiveness of the laboratory’s pre-examination, examination and post-examination processes and quality management system. | 100
|  | Service Agreements How the laboratory will: 1) establish service agreements; 2) review service agreements; 3) handle walk in patients (if applicable); 4) inform customers and users of any changes that affect the results of the requisition stated on the service agreement; 5) communicate to the requester of any work that has been referred; 6) retain records of communication? |  |  |  |  | ISO15189:2012 Clause 4.4.1 and 5.4 Notes: By accepting a requisition form from an authorized requester, the laboratory is considered to have entered into a service agreement. Customers and users may include clinicians, health care organizations, third party payment organizations or agencies, pharmaceutical companies, and patients. | 100
|  | Examination by Referral Laboratories and Consultants How the laboratory will: 1) select referral laboratories and consultants who provide opinions as well as interpretations; 2) evaluate and monitor the performance of referral laboratories and consultants who provide opinions as well as interpretations; 3) maintain a list of approved referral laboratories and consultants; 4) maintain a records of referred samples; 5) tracking of referred samples and their results; 6) report results from referral labs; 7) package and transport referred samples; 8) record communication of results from referral laboratories and consultants? |  |  |  |  | ISO15189:2012 Clause 4.5 and 5.8 and 4.13 Note: The laboratory must have a documented procedure for selecting and evaluating referral laboratories and consultants who provide opinions as well as interpretation for complex testing in any discipline. | 100
|  | External Services and Suppliers How the laboratory will: 1) select external purchases and services; 2) establish its selection criteria, including acceptance and rejection criteria; 3) approve and maintain its approved suppliers list; 4) define the requirements of its purchase supplies and services; 5) review and monitor the performance of its approved suppliers; 6) establish frequency of reviews? |  |  |  |  | ISO15189:2012 Clause 4.6 and 5.3 Note: The laboratory must have a documented procedure for the selection and purchasing of external services, equipment, reagents and consumable supplies that affect the quality of its service. | 100
|  | Purchasing and Inventory Control How the laboratory will: 1) request, order and receive supplies; 2) establish acceptance/rejection criteria for purchased items; 3) store purchased supplies; 4) control their inventory; 5) monitor and handle expired consumables? |  |  |  |  | ISO15189 :2012 Clause 4.6 and 5.3.2 Note: The laboratory shall have a documented procedure for the reception, storage, acceptance testing and inventory management of reagents and consumables. | 100
|  | Advisory Services How the laboratory will: 1) advise on the choice of examinations it offers; 2) communicate its advisory services to its users; 3) advise on clinical indications and limitations of examination procedures; 4) advise on the frequency of examination; 5 ) provide individual clinical case advice; 6) advise on interpretation of results; 7) promote the effective utilization of laboratory services; 8) provide consultation on scientific and logistic matters; 9) advise on the required type of sample and volume for testing? |  |  |  |  | ISO15189:2012 Clause 4.7 Note: The laboratory must have a system in place for providing advise to its users. | 100
|  | Resolution of Complaints and Feedback How the laboratory will: 1) manage complaints received from clinicians, patients, laboratory staff or other parties; 2) collect, receive and handle feedback received from clinicians, patients, laboratory staff or other parties; 3) keep records of all complaints, the investigations and actions taken, 4) determine the timeframe for closure and feedback to the complainant; 5) monitor effectiveness of corrective and preventative actions taken on complaints and feedback? |  |  |  |  | ISO15189:2012 Clause 4.8 and 4.10 Note: The laboratory must have a documented procedure for the management of complaints or other feedback received from clinicians, patients, laboratory staff or other parties. Records shall be maintained of all complaints and their investigation and the action taken | 100
|  | Identification and Control of Nonconformities (NC) How the laboratory will: 1) identify types of nonconformities in any aspect of the quality management system from pre, analytic and post analytic; 2) record NCs (how and where); 3) assign who is responsible for resolving the NC; 4) determine time frame for resolving NCs; 5) halt examinations (by an authorized person); 6) ensure the recall of released results of nonconforming or potentially nonconforming examinations; 7) release results after corrective action has been taken? |  |  |  |  | ISO15189:2012 Clause 4.9 Note: Nonconforming examinations or activities occur in many different areas and can be identified in many different ways, including clinician complaints, internal quality control indications, and instrument calibrations, checking of consumable materials, inter-laboratory comparisons, staff comments, reporting and certificate checking, laboratory management reviews, and internal and external audits. | 100
|  | Corrective Action (CA) How the laboratory will: 1) determine the root cause; 2) evaluate the need for CA to ensure that NCs do not recur; 3) assign the person responsible for the CA; 4) determine and implement CA(including person responsible and timeframe); 5) record CA taken; 6) monitor and review the effectiveness of the CA taken? |  |  |  |  | ISO15189:2012 Clause 4.10 Note: Action taken at the time of the nonconformity to mitigate effects is considered “immediate” action. Only action taken to remove the root cause of the problem that is causing the Non Conformities is considered “corrective” action. Any immediate action taken must also be documented. Corrective actions must be appropriate to the effects of the nonconformities encountered. | 100
|  | Preventive Action (PA) How the laboratory will: 1) review laboratory data and information to determine potential nonconformities; 2) determine the root cause(s) of potential nonconformities; 3) evaluate the need for preventive action; 4) record the PA; 5) determine and implement PA (including person responsible and timeframe); 6) monitor and review the effectiveness of implementation of PA? |  |  |  |  | ISO15189 :2012 Clause 4.11 Note: Preventive action is a proactive process for identifying opportunities for improvement rather than a reaction to the identification of problems or complaints (i.e. nonconformities). In addition to review of the operational procedures, preventive action might involve analysis of data, including trend and risk analyses and external quality assessment (proficiency testing). The laboratory shall determine action to eliminate the causes of potential nonconformities in order to prevent their occurrence. Preventive actions shall be appropriate to the effects of the potential problems. | 100
|  | Continual Improvement How the laboratory will: 1) identify improvement activities within the Quality Management System; 2) develop improvement plans; 3) record improvement plans; 4) implement action plans; 5) communicate improvement plans and related goals to staff? |  |  |  |  | ISO15189:2012 Clause 4.1.1.2; 4.12; 4.14.5 Note: Improvement activities must be identified within the pre-examination, examination and post-examination processes. Laboratory management shall ensure that the laboratory participates in continual improvement activities that encompass relevant areas and outcomes of patient care. | 100
|  | Internal Audits How the laboratory will: 1) determine an audit schedule; 2) determine the roles and responsibilities for planning and conducting audits; 3) select the auditors; 4) define the types of audits; 4) define the frequency of audits; 5) define the scope of the internal audit; 6) record audit findings (forms and reports); 7) ensure corrective action is taken for all nonconformities identified with in the allocated time frame; 8) closure of Non Conformities identified during audits? |  |  |  |  | ISO15189:2012 Clause 4.14.5 Note: The cycle for internal auditing should normally be completed in one year. It is not necessary that internal audits cover each year, in depth, all elements of the quality management system. The laboratory may decide to focus on a particular activity without completely neglecting the others. The laboratory shall conduct internal audits at planned intervals to determine whether all activities in the quality management system, including pre-examination, examination, and post-examination. | 100
|  | Risk Management How the laboratory will: 1) evaluate the impact of potential pitfalls on work processes and examination results that affect patient results? (Refer to Question 6.3 of this checklist) |  |  |  |  | ISO15189:2012 Clause 4.14.6 Notes: Risk must be managed at the pre-examination processes, examination processes and post examination processes. The laboratory shall evaluate the impact of work processes and potential failures on examination results as they affect patient safety, and shall modify processes to reduce or eliminate the identified risks and document decisions and actions taken. | 100
|  | Management Review How the laboratory will: 1) define frequency of having management reviews; 2) define the agenda (input); 3) determine the key attendees; 4) record decisions and actions to be taken (output); 5) assign a person responsible and due dates for actions arising; 6) communicate decisions and actions to be taken to the relevant persons including laboratory staff; 7) ensure all actions arising are completed within the defined timeframe? (refer to Question 2.2 of this checklist for the agenda of the meeting) |  |  |  |  | ISO15189:2012 Clause 4.15 Note: Laboratory management shall review the quality management system at planned intervals to ensure its continuing suitability, adequacy and effectiveness and support of patient care. | 100
|  | Personnel Management How the laboratory will: 1) define the structure of the organization (organizational plan); 2) manage personnel (personnel policies); 3) maintain personnel records? (refer to Question 3.5 of the checklist for list of personnel records required) |  |  |  |  | ISO15189:2012 Clause 5.1.1; 5.1.9; 4.13 Note: The laboratory must have a documented procedure for personnel management and maintain records for all personnel to indicate compliance with requirements. | 100
|  | Personnel Training How the laboratory will: 1) perform staff orientation; 2) conduct initial and refresher training; 3) provide a continuous education program; 4) identify required training relevant to job title and responsibilities; 5) keep record of training; 6) evaluate the effectiveness of training? |  |  |  |  | ISO15189:2012 Clause 4.1.1.4 and 5.1.5 Note: Training includes external and internal trainings. The effectiveness of the training programme must be periodically reviewed. | 100
|  | Competency Assessment How the laboratory will: 1) assess the competence of personnel to perform assigned managerial or technical tasks; 2) assess ongoing competency; 3) establish competency criteria; 4) provide feedback to persons assessed; 5) schedule retraining based on the assessment outcome; 6) keep records of competency assessments and outcomes? |  |  |  |  | ISO15189 :2012 Clause 4.1.1.4 and 4.4 and 5.1.6 Note: Competency could be assessed using a combination of some or all of the following methods: direct observation; monitoring and the recording of examination results; review of work records; problem solving skills; blinded samples, review of accumulative IQC and EQA. Competency assessment for professional judgment should be designed as specific and fit for purpose. | 100
|  | Authorization How the laboratory will: 1) document authorization levels for the different tasks and roles; 2) appoint deputies for the key positions where appropriate? |  |  |  |  | ISO15189:2012 Clause 4.1.2 Note: Authorization may be in the form of a Job description, letter of appointment, approved authority matrix etc. | 100
|  | Review of Staff Performance How the laboratory will: 1) plan and perform staff appraisals; 2) establish frequency of monitoring and review of staff performance outcome; 3) keep records of staff performance; 4) train staff who perform staff appraisals? |  |  |  |  | ISO15189:2012 Clause 4.1.2.1 and 5.1.7 Note: In addition to the assessment of technical competence, the laboratory management must ensure that reviews of staff performance consider the needs of the laboratory and of the individual in order to maintain or improve the quality of service given to the users and encourage productive working relationships. Staff performing reviews should receive appropriate training. | 100
|  | Accommodation and Environmental Conditions How the laboratory will: 1) evaluate and determine the sufficiency and adequacy of the space allocated for the performance of and scope of work; 2) ensure the laboratory and office facilities are suitable for the tasks to be undertaken; 3) ensure the storage and disposal facilities meet the applicable requirements; 4) ensure staff have space for staff activities (supply of drinking water, storage space for personal and protective equipment and clothing); 5) monitor, control and record any specific environmental and accommodation requirements? |  |  |  |  | ISO15189:2012 Clause 4.1.1.4 and 5.2; 5.2.6 Note: The laboratory must have space allocated for the performance of its work that is designed to ensure the quality, safety and efficacy of the service provided to the users and the health and safety of laboratory personnel, patients and visitors. The laboratory shall evaluate and determine the sufficiency and adequacy of the space allocated for the performance of the work. Evaluating and determining the sufficiency and adequacy of space may be done during internal audits, risk assessments or at management review meeting, however it must be documented that it was evaluated and found to be adequate. | 100
|  | Laboratory equipment How the laboratory will: 1) select equipment; 2) purchase equipment; 3) manage equipment; 4) maintain equipment records 5) capture the minimum information on equipment label; 6) manage defective equipment; 7) define the equipment maintenance frequency; 8) record the maintenance ; 9) prevent unauthorized use (access control) of equipment; 10) manage obsolete equipment; 11) manage safe handling, transportation, storage and use to avoid deterioration and contamination, 12) track and verify completion of repairs? |  |  |  |  | ISO15189:2012 Clause 4.13; 5.3.1.1; 5.3.1.3 NOTE: For the purposes of this checklist, laboratory equipment includes hardware and software of instruments, measuring systems, and laboratory information systems. The laboratory shall have a documented procedure for the selection, purchasing and management of equipment. | 100
|  | Calibration of Equipment How the laboratory will: 1) define frequency of calibration; 2) handle in house calibrations (pipettes, thermometers, timers etc.); 3) record calibration status (use of stickers and calibration certificates); 4) handle failed calibrations? |  |  |  |  | ISO15189:2012 Clause 5.3.1.4 Notes: The laboratory must have a documented procedure for the calibration of equipment that directly or indirectly affects examination results. Documentation of calibration traceability to a higher order reference material or reference procedure may be provided by an examination system manufacturer. Such documentation is acceptable as long as the manufacturer's examination system and calibration procedures are used without modification. | 100
|  | Pre-examination Processes How the laboratory will provide information for patients and users on: 1) primary sample collection and handling; 2) instructions for pre-collection activities; 3) instructions for collection activities; 4) preparation and storage prior to dispatch to the laboratory; 5) sample and volume requirements; 6) Sample transportation; 7) time limits and special handling; 8) acceptance and rejection criteria; 9) confidentiality; 10) complaints procedure? |  |  |  |  | ISO15189:2012 Clause 5.4; 5.4.1; 5.4.3; 5.4.4.1; 5.4.5; 5.4.6; 5.4.7 Note: The laboratory must have documented procedures and information for pre-examination activities to ensure the validity of the results of examinations. | 100
|  | Validation and Verification of examination procedures / Equipment How the laboratory will: 1) select testing procedures; 2) perform equipment validation; 3) perform method validation; 4) perform equipment verification; 5) perform method verification; 6) define validation /verification protocol specific for each procedure at the time of validation or verification; 7) compare results from the different procedures, equipment, methods being used for the same test either located at the same site or at different sites? |  |  |  |  | ISO15189:2012 Clause 5.5.1.2; 5.6.4 and 5.5.1.3 Note: Validations should be done on a) non-standard methods; b) laboratory designed or developed methods; c) standard methods used outside their intended scope; d) validated methods subsequently modified. "Verification” is performed on methods that are being used without any modifications and is a process of evaluating of whether or not the procedure meets the performance characteristics stated by the manufacturer i.e. the manufacturer validation claims. The performance characteristics are obtained from the manufacture (validation reports) or from package inserts. Comparison of different methods used for same tests is ongoing verification. The frequency and characteristics to be checked in ongoing verification must be clearly defined. Note: All procedures or equipment used as backup must also be validated/verified as relevant. | 100
|  | Measurement of Uncertainty How the laboratory will: 1) determine Measurement of uncertainty on measured quantity values (quantitative tests); 2) define the performance requirements for the measurement uncertainty (e.g Standard Deviation; Clinical decision points)? Refer to Question 5.4 on this checklist |  |  |  |  | ISO15189:2012 Clause 5.5.1.4 Note: Uncertainty of measurement is used to indicate the confidence we have that the reported figure is correct. Uncertainty of measurement may be calculated using the calculated CV of at least 30 sets of internal QC data: CV% x 2 = Uncertainty of measurement (UM). The laboratory shall calculate the UM for all quantitative tests. These shall only be reported to clinicians if they request for them. For well-established methods, it is recommended a minimum of six months internal QC data should be used to calculate UM, updated at least annually where possible. For new methods at least 30 data points for each level of QC across at least two different batches of calibrator and reagents should be used to provide an interim estimate of uncertainty of measurement. | 100
|  | Biological Reference Intervals or Clinical Decision Values How the laboratory will: 1) define the biological reference intervals; 2) document the source of the reference intervals; 3) communicate changes to the users? |  |  |  |  | ISO15189:2012 Clause 5.5.2 Note: The laboratory shall define the biological reference intervals or clinical decision values, document the basis for the reference intervals or decision values and communicate this information to users. | 100
|  | Documentation of examination procedures How the laboratory will: 1) format general and technical Standard Operating Procedures; 2) define the minimum requirements for a SOP? |  |  |  |  | ISO15189:2012 Clause 5.5.3 Note: Working instructions, card files or similar systems that summarize key information are acceptable for use as a quick reference at the workbench, provided that a fully documented procedure is available for reference. Information from product instructions for use may be incorporated into examination procedures by reference in the SOP. The minimum requirements for a technical SOP should be a) purpose of the examination; b) principle and method of the procedure used for examinations; c) type of sample; d) required equipment and reagents; e) environmental and safety controls; f) procedural steps; g) interferences (e.g. lipemia, hemolysis, bilirubinemia, drugs) and cross reactions; h) principle of procedure for calculating results; i) laboratory clinical interpretation; j) potential sources of variation; k) references. | 100
|  | Laboratory Contingency Plan How the laboratory will ensure that there are no interruption to services in the event of: 1) staff shortage; 2) equipment breakdown; 3) prolonged power outages; 4) stock outs of reagents and consumables; 5) fire, natural disasters e.g. severe weather or floods, bomb threat or civil disturbances; 7) LIS failure? |  |  |  |  | ISO15189:2012 Clause 4.1.1.4; 5.2; 5.3.1; 5.10 Notes: the laboratory should maintain sufficient replacement parts to minimize testing downtime (e.g. pipette components, microscope bulbs and fuses, safety caps or buckets for safety centrifuge). Contingency plans should be periodically tested. Where the laboratory uses another laboratory as a backup, the performance of the back-up laboratory shall be regularly reviewed to ensure quality results | 100
|  | Quality control and Quality Assurance How the laboratory will: 1) use IQC and EQA (Interlaboratory comparison); 2) define the frequency of processing IQC; 3) define the acceptable ranges; 4) Evaluate and monitor laboratory performance using EQA and QC data; 5) troubleshoot unacceptable EQA and QC; 6) compare results using different procedures, equipment and sites; 7) notify users of any differences in comparability of results? |  |  |  |  | ISO15189:2012 Clause 4.10; 5.6; 5.6.2.1; 5.6.2.3; 5.6.3.1 Note: The laboratory should choose concentrations of control materials, wherever possible, especially at or near clinical decision values, which ensure the validity of decisions made. Use of independent third party control materials should be considered, either instead of, or in addition to, any control materials supplied by the reagent or instrument manufacturer. EQA should cover the pre-examination process, examination process and post examination process. Where an EQA program is not available, the laboratory can use alternative methods with clearly defined acceptable results e.g. exchange of samples with other labs, testing certified materials, sample previously tested. All procedures or equipment used as backup must also be included in EQA programme. | 100
|  | Reporting and Release of Results How the laboratory will : 1) issue standardized report (define the format and medium); 2) review patient results; 3) communicate patient results including alert, urgent and critical results; 4) ensure release of results to authorized persons; 5) amend reports; 6) issue of amended reports; 7) store patient results; 8) maintain patient results. (Refer to Question 9.3 of this checklist) |  |  |  |  | ISO15189:2012 Clause 5.8.1; 5.9.1 Note: Reports may be issued as a hard copy or electronically, all results issued verbally must be followed by a final report. The results of each examination must be reported accurately, clearly, unambiguously and in accordance with any specific instructions in the examination procedures. The laboratory must define the format and medium of the report (i.e. electronic or paper) and the manner in which it is to be communicated from the laboratory. | 100
|  | Laboratory Information System (LIS) (Computerized or non-computerized) How the laboratory will: 1) select a LIS; 2) verify /validate the LIS; 3) define authorities and responsibilities for the management and use of the information system; 4) ensure patient confidentiality is maintained at all times; 5) maintain the system; 6) back-up data; 7) safeguard against tempering by un-authorized users? |  |  |  |  | ISO15189:2012 Clause 5.10 Note: “information systems” includes the management of data and information contained in both computer and non-computerized systems. Some of the requirements may be more applicable to computer systems than to non-computerized systems. Computerized systems can include those integral to the functioning of laboratory equipment and stand-alone systems using generic software, such as word processing, spreadsheet and database applications that generate, collate, report and archive patient information and reports. | 100
|  | Laboratory Safety Manual How the laboratory will: 1) ensure all safety measures are implemented at the laboratory as applicable to national and international guidelines and regulations? (Refer to section 12 of this checklist for the contents of a safety manual) |  |  |  |  | ISO15190:2013 Clause 4.1.1.4; 5.2 Note: Laboratory management must implement a safe laboratory environment in compliance with good practice and applicable requirements. | 100
|  | Answer to AMR question G1.1 | [5 empty] | 100
|  | Answer to AMR question G1.2 | [5 empty] | 100
|  | Answer to AMR question G1.3 | [5 empty] | 100
|  | Answer to AMR question G1.4 | [5 empty] | 100
|  | Answer to AMR question U1.1 | [5 empty] | 100
|  | Answer to AMR question U1.2 | [5 empty] | 100
|  | Answer to AMR question U1.3 | [5 empty] | 100
|  | Answer to AMR question F1.1 | [5 empty] | 100
|  | Answer to AMR question F1.2 | [5 empty] | 100
|  | Answer to AMR question F1.3 | [5 empty] | 100
|  | Answer to AMR question B1.1 | [5 empty] | 100
|  | Answer to AMR question B1.2 | [5 empty] | 100
|  | Answer to AMR question B1.3 | [5 empty] | 100
|  | Answer to AMR question N1.1 | [5 empty] | 100
|  | Answer to AMR question N1.2 | [5 empty] | 100
|  | Answer to AMR question N1.3 | [5 empty] | 100
|  | Answer to AMR question P1.1 | [5 empty] | 100
|  | Answer to AMR question P1.2 | [5 empty] | 100
|  | Answer to AMR question P1.3 | [5 empty] | 100
|  | Answer to AMR question W1.1 | [5 empty] | 100
|  | Answer to AMR question W1.2 | [5 empty] | 100
|  | Answer to AMR question W1.3 | [5 empty] | 100 | 5700
1.6 | Policy and SOPs Accessibility Are policies and SOPs easily accessible/ available to all staff and written in a language commonly understood by respective staff? |  | 2 |  | 0 |  | ISO15189:2012 Clause 4.2.2.1; 4.3; 5.5 Note: All documentation must be current and approved by an authorized person. The documentation can be in any form or type of medium, providing it is readily accessible and protected from unauthorized changes and undue deterioration
1.7 | Policies and SOPs Communication Is there documented evidence that all relevant policies and SOPs have been communicated to and are understood and implemented by all staff as related to their responsibilities? |  | 2 |  | 0 |  | ISO15189:2012 Clause 4.2.2.2; 5.1.5(b) Note: The lab must have a system in place to ensure all staff are aware of the contents of all documents. All laboratory staff shall have access to and be instructed on the use and application of the quality manual and the referenced documents.
1.8 | Document Control Log Are policies and procedures dated to reflect when it was put into effect, its location, when it was reviewed and when it was discontinued? |  | 2 |  | 0 |  | ISO15189:2012 Clause 4.3 Note: Current authorized editions and their distribution are identified by means of a list (e.g. document register, log or master index).
1.9 | Discontinued Policies and SOPs Are invalid or discontinued policies and procedures clearly marked / identified and removed from use and one copy retained for reference purposes? |  | 2 |  | 0 |  | ISO15189:2012 Clause 4.3 Note: Obsolete controlled documents are dated and marked as obsolete. At least one copy of an obsolete controlled document is retained for a specified time period or in accordance with applicable specified requirements.
1.10 | Data Files Are test results, technical and quality records, invalid or discontinued policies and procedures archived for a specified time period in accordance with national/international guidelines? |  | 2 |  | 0 |  | ISO15189:2012 Clause 4.3; 4.13 Note: Copies or files of results should be archived. The retention period may vary; however, the reported results shall be retrievable for as long as medically relevant or as required by national, regional or local authorities.
1.11 | Archived Results Accessibility Is there an archiving system that allows for easy and timely retrieval of archived records and results? |  | 2 |  | 0 |  | ISO15189:2012 Clause 4.13 Note: Records can be in any form or type of medium providing they are readily accessible and protected from unauthorized alterations. Archived patient results must be easily, readily and completely retrievable within a timeframe consistent with patient care needs.
| Totals |  | 28 |  | 0
Yes
Partial
No
N/A

## SLIPTA_S2
Section 2 - Management Reviews
Question no. | Question |  | Available Score | Answer | Score | Comments | Standard
2.1 | Routine Review of Quality and Technical Records Does the laboratory supervisor routinely perform a documented review of all quality and technical records? |  | 5 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 4.1.1.4; 4.2.1 There must be documentation that the laboratory manager/supervisor or a designee reviews the quality programme regularly. The review must ensure that recurrent problems have been addressed, and that new or redesigned activities have been evaluated.
Does the supervisor’s review include the following?
|  | a) Follow-up of action items from previous reviews | [5 empty] | 100
|  | b) Status of corrective actions taken and required preventive actions | [5 empty] | 100
|  | c) Reports from personnel | [5 empty] | 100
|  | d) Environmental monitoring log sheets | [5 empty] | 100
|  | e) Specimen rejection logbook | [5 empty] | 100
|  | f) Equipment calibration and maintenance records | [5 empty] | 100
|  | g) IQC records across all test areas | [5 empty] | 100
|  | h) Outcomes of PTs and other forms of inter-laboratory comparisons | [5 empty] | 100
|  | i) Quality indicators | [5 empty] | 100
|  | j) Customer complaints and feedback | [5 empty] | 100
|  | k) Results of improvement projects | [5 empty] | 100
|  | l) Documentation of review and action planning with staff for resolution and follow-up review | [5 empty] | 100 | 1200
2.2 | Management Review Does the laboratory management annually perform a review of all quality systems at a management review meeting? |  | 5 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 4.1.1.4; 4.15.2; 4.15.4 Note: The interval between management reviews should be no greater than 12 months; however, shorter intervals should be adopted when a quality management system is being established.
Does the management review meeting include the following INPUTS?
|  | a) The periodic review of requests, and suitability of procedures and sample requirements | [5 empty] | 100
|  | b) Assessment of user feedback | [5 empty] | 100
|  | c) Staff suggestions | [5 empty] | 100
|  | d) Internal audits | [5 empty] | 100
|  | e) Risk management | [5 empty] | 100
|  | f) Use of quality indicators | [5 empty] | 100
|  | g) Assessments by external organizations | [5 empty] | 100
|  | h) Results of participation in interlaboratory comparison programmes (PT/EQA) | [5 empty] | 100
|  | i) Monitoring and resolution of complaints | [5 empty] | 100
|  | j) Performance of suppliers | [5 empty] | 100
|  | k) Identification and control of nonconformities | [5 empty] | 100
|  | l) Results of continual improvement including, current status of corrective actions and preventive actions | [5 empty] | 100
|  | m) Follow-up actions from previous management reviews | [5 empty] | 100
|  | n) Changes in the volume and scope of work, personnel, and premises that could affect the quality management system | [5 empty] | 100
|  | o) Recommendations for improvement, including technical requirements | [5 empty] | 100
|  | p) Review of quality objectives and the quality policy for appropriateness and continuous improvement | [5 empty] | 100
Does the management review meeting include the following OUTPUTS?
|  | a) Are management review outputs recorded? | [5 empty] | 100
|  | b) Does the output records of the MR meeting capture decisions made, persons responsible for actions to be taken and timeframes? | [5 empty] | 100
|  | c) Does the report address resources required (human, financial, material)? | [5 empty] | 100
|  | d) Does it refer to improvement for the users? | [5 empty] | 100
|  | e) Does it refer to improvement of the effectiveness of the quality system? | [5 empty] | 100
|  | f) Were the quality objectives and the quality policy reviewed for appropriateness and continuous improvement? | [5 empty] | 100
|  | Answer to AMR question G2.1 | [5 empty] | 100
|  | Answer to AMR question G2.2 | [5 empty] | 100
|  | Answer to AMR question G2.3 | [5 empty] | 100
|  | Answer to AMR question G2.4 | [5 empty] | 100 | 2600
2.3 | Are findings and actions from MR communicated to the relevant staff? |  | 2 |  | 0 |  | ISO15189:2012 Clause 4.1.1.4; 4.15.4 Note: Findings and actions arising from management reviews shall be recorded and reported to laboratory staff.
2.4 | Does lab management ensure actions from MR are completed within defined timeframes? |  | 2 |  | 0 |  | ISO15189:2012 Clause 4.1.1.4; 4.15.4 Note: Laboratory management shall ensure that actions arising from management review are completed within a defined timeframe.
| Totals |  | 14 |  | 0
Yes
Partial
No
N/A

## SLIPTA_S3
Section 3 - Organization & Personnel
Question no. | Question |  | Available Score | Answer | Score | Comments | Standard
3.1 | Duty Roster and Daily Routine Does the laboratory have a duty roster that covers normal and after hours? |  | 2 |  | 0 |  | ISO15189:2012 Clause 4.1.1.4(c); 4.1.2.1(i) Note: A duty roster designates specific laboratory personnel to specific workstations. Daily routines should be prioritized, organized and coordinated to achieve optimal service delivery for patients.
3.2 | Organizational Chart and External/Internal Reporting Systems Is an organizational chart available that indicates the relationship between the laboratory and its parent organization? |  | 2 |  | 0 |  | ISO15189:2012 Clause 4.1.2.5 Note: An up-to-date organizational chart and/or narrative description should be available detailing the external and internal reporting relationships for laboratory personnel. The organizational chart or narrative should clearly show how the laboratory is linked to the rest of the hospital and laboratory services where applicable.
3.3 | Laboratory Director Is the laboratory directed by a person(s) with the competency and delegated responsibility to perform the following: |  | 3 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 4.1.1.4 Note: a director may be a person(s) with responsibility for, and authority over, a laboratory. The person or persons referred to may be designated collectively as laboratory director. Other settings may not use the term “Lab Director” but in this question, it refers to person/persons that are running the laboratory, however they decide to name them
|  | a) Provide effective leadership, budgeting and planning | [5 empty] | 100
|  | b) Communicate with stakeholders | [5 empty] | 100
|  | c) Ensure adequate competent staff | [5 empty] | 100
|  | d) Ensure the implementation of the QMS | [5 empty] | 100
|  | e) Selection and monitoring of lab supplies | [5 empty] | 100
|  | f) Selection and monitoring of referral labs | [5 empty] | 100
|  | g) Ensure a safe lab environment | [5 empty] | 100
|  | h) Advisory services | [5 empty] | 100
|  | i) Provide professional development programs for laboratory staff | [5 empty] | 100
|  | j) Address complaints, requests or suggestions from staff and/or lab users | [5 empty] | 100
|  | k) Design and implement a contingency plan | [5 empty] | 100 | 1100
3.4 | Quality Management System Oversight Is there a quality officer/manager with delegated responsibility to oversee compliance with the quality management system? |  | 3 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 4.1.2.7 There should be a quality manager (however named) with delegated authority to oversee compliance with the requirements of the quality management system. The quality manager must report directly to the level of laboratory management at which decisions are made on laboratory policy and resources.
|  | a) Is there an appointment letter, job description or terms of reference available? | [5 empty] | 100
|  | b) Does the quality manager ensure that processes needed for the quality management system are established, implemented, and maintained? | [5 empty] | 100
|  | c) Does the QM report to management which decisions relating to quality are made? | [5 empty] | 100
|  | d) Does the QM promote awareness of users’ needs and requirements throughout the organization? | [5 empty] | 100
|  | e) Does the QM participate in management reviews? | [5 empty] | 100 | 500
3.5 | Personnel Filing System Are records of personnel maintained and do they include the following? |  | 3 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 5.1.9 Note: Personnel files must be maintained for all current staff. Wherever (offsite or onsite) and however the records are kept, the records must be easily accessible when required. In some laboratories, not all records may be kept in a single file in one place e.g. training and competency records should be kept in the laboratory, medically related information with the administration.
If files are present, do they document or contain the following:
|  | a) Educational and professional qualifications | [5 empty] | 100
|  | b) Copy of certification or license to practice, when applicable | [5 empty] | 100
|  | c) Previous work experience e.g. CV | [5 empty] | 100
|  | d) Job description | [5 empty] | 100
|  | e) Record of introduction of new staff to the laboratory environment | [5 empty] | 100
|  | f) Training in current job tasks including vendor training received on-site | [5 empty] | 100
|  | g) Records of competency assessments | [5 empty] | 100
|  | h) Records of continuing education | [5 empty] | 100
|  | i) Reviews of staff performance | [5 empty] | 100
|  | j) Reports of accidents and exposure to occupational hazards | [5 empty] | 100
|  | k) Immunization status, as applicable, relevant to assigned duties | [5 empty] | 100
|  | l) Letter of employment or appointment | [5 empty] | 100
|  | m) Employee medical surveillance records | [5 empty] | 100 | 1300
3.6 | Laboratory Staff Training Is there a system for training that covers the following? |  | 3 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 4.1.1.4(c); 5.1.5 Note: The effectiveness of the training program shall be reviewed regularly. Personnel that are undergoing training shall be supervised at all times.
|  | a) The quality management system | [5 empty] | 100
|  | b) Assigned work processes, procedures and tasks | [5 empty] | 100
|  | c) The applicable laboratory information system | [5 empty] | 100
|  | d) Health and safety, including the prevention or containment of the effects of adverse incidents | [5 empty] | 100
|  | e) Laboratory Ethics | [5 empty] | 100
|  | f) Confidentiality of patient information | [5 empty] | 100
|  | g) Is there supervision for persons undergoing training | [5 empty] | 100
|  | h) Continuous medical education | [5 empty] | 100
|  | i) Review of effectiveness of the training program | [5 empty] | 100
|  | Answer to AMR question U3.1 | [5 empty] | 100
|  | Answer to AMR question F3.1 | [5 empty] | 100
|  | Answer to AMR question B3.1 | [5 empty] | 100
|  | Answer to AMR question N3.1 | [5 empty] | 100
|  | Answer to AMR question P3.1 | [5 empty] | 100
|  | Answer to AMR question W3.1 | [5 empty] | 100 | 1500
3.7 | Staff Competency Assessment and Retraining Is there a system for competency assessment that covers the following? |  | 3 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 4.1.2.1(h); 5.1.6 Note: Newly hired lab staff must be assessed for competency before performing independent duties. All lab staff must undergo ongoing competency at a frequency defined by the laboratory. Staff assigned to a new section should be assessed before fully assuming independent duties. When deficiencies are noted, retraining and reassessment must be planned and documented. If the employee’s competency remains below standard, further action might include supervisory review of work, re-assignment of duties, or other appropriate actions. Records of competency assessments and resulting actions should be retained in personnel files and/or quality records. Records should show which skills were assessed, how those skills were measured, and who performed the assessment.
|  | a) Are competency assessments performed according defined criteria | [5 empty] | 100
|  | b) New hires | [5 empty] | 100
|  | c) Existing staff | [5 empty] | 100
|  | d) Retraining and re-assessment where needed | [5 empty] | 100
|  | Answer to AMR question U3.2 | [5 empty] | 100
|  | Answer to AMR question F3.2 | [5 empty] | 100
|  | Answer to AMR question B3.2 | [5 empty] | 100
|  | Answer to AMR question N3.2 | [5 empty] | 100
|  | Answer to AMR question P3.2 | [5 empty] | 100
|  | Answer to AMR question W3.2 | [5 empty] | 100 | 1000
3.8 | Staff meetings Are staff meetings held regularly? |  | 3 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 4.1.2.1(a); (e); 4.1.2.2; 4.1.2.6; 4.4; 4.14.3 Note: The laboratory should hold regular staff meetings to ensure communication within the laboratory. Meetings should have recorded notes to facilitate review of progress over time.
Do meetings include the following items?
|  | a) Follow-up of action items from previous staff meetings | [5 empty] | 100
|  | b) Systemic and/or recurrent problems and issues addressed, including actions to prevent recurrence | [5 empty] | 100
|  | c) Complaints | [5 empty] | 100
|  | d) Communication on reviewed/revised/redundant SOPs | [5 empty] | 100
|  | e) Review of results from prior corrective actions | [5 empty] | 100
|  | f) Discussion and evaluation of improvement topics/projects | [5 empty] | 100
|  | g) Feedback given by staff that have attended hospital meetings, external meetings, training, conferences, workshops, etc. | [5 empty] | 100
|  | h) Relay of reports and updates from lab attendance at meetings with clinicians (the use of lab services and/or attendance at clinical rounds) | [5 empty] | 100
|  | i) Recording and monitoring of meeting notes for progress on issues. | [5 empty] | 100 | 900
| Totals |  | 22 |  | 0
Yes
Partial
No
N/A

## SLIPTA_S4
Section 4 - Client Management and Customer Service
Question no. | Question |  | Available Score | Answer | Score | Comments | Standard
4.1 | Advice and Training by Qualified Staff |  | 2 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 4.1.1.4(g); 4.7 Note: Authorized staff should provide advice on sample type, examination choice, frequency and results interpretation.
|  | Do staff members with appropriate professional qualifications provide clients with advice and/or training regarding required types of samples, choice of examinations, repeat frequency, and interpretation of results? | [5 empty] | 100
|  | Answer to AMR question U4.1 | [5 empty] | 100
|  | Answer to AMR question U4.2 | [5 empty] | 100
|  | Answer to AMR question F4.1 | [5 empty] | 100
|  | Answer to AMR question F4.2 | [5 empty] | 100
|  | Answer to AMR question B4.1 | [5 empty] | 100
|  | Answer to AMR question B4.2 | [5 empty] | 100
|  | Answer to AMR question N4.1 | [5 empty] | 100
|  | Answer to AMR question N4.2 | [5 empty] | 100
|  | Answer to AMR question N8.1 | [5 empty] | 100
|  | Answer to AMR question N8.2 | [5 empty] | 100
|  | Answer to AMR question N8.3 | [5 empty] | 100
|  | Answer to AMR question N8.4 | [5 empty] | 100
|  | Answer to AMR question N8.5 | [5 empty] | 100
|  | Answer to AMR question N8.6 | [5 empty] | 100
|  | Answer to AMR question P4.1 | [5 empty] | 100
|  | Answer to AMR question P4.2 | [5 empty] | 100
|  | Answer to AMR question W4.1 | [5 empty] | 100
|  | Answer to AMR question W4.2 | [5 empty] | 100 | 1900
4.2 | Resolution of Complaints Does the laboratory investigate (review) and resolves of customer complaints? |  | 2 |  | 0 |  | ISO15189:2012 Clause 4.1.1.4(m); 4.8; 4.15.2(i) Note: The laboratory must have a documented procedure for the management of complaints or other feedback received from clinicians, patients, laboratory staff or other parties. Feedback must be given to the complainant.
4.3 | Laboratory Handbook for Clients - information for users |  | 2 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 4.1.1.4(g); 4.5; 5.4.2 Note: The laboratory should provide its clients with a handbook that outlines the laboratory’s hours of operation, available tests, specimen collection instructions, packaging and shipping directions, and expected turnaround times.
|  | Is there a laboratory handbook for laboratory users that includes information on: - location of the laboratory, - services offered, - laboratory operating times, - instructions on completion of request forms, - instruction for preparation of the patient; - sample collection including patient collected samples, - transport, - agreed turnaround times, - acceptance and rejection criteria, - availability of advice on examination and interpretation of results; - lab policy on protection of personal information, - laboratory complaints procedure. | [5 empty] | 100
|  | Answer to AMR question G4.1 | [5 empty] | 100 | 200
4.4 | Communication Policy on Delays in Service Is timely, documented notification provided to customers when the laboratory experiences delays or interruptions in testing (due to equipment failure, stock-outs, staff levels, etc.) or finds it necessary to change examination procedures, and when testing resumes? |  | 2 |  | 0 |  | ISO15189:2012 Clause 4.1.2.6; 4.4; 5.8.1 There should be a policy for notifying the requester when an examination is delayed. Such notification shall be documented for both service interruption and resumption as well as related feedback from clinicians. Clinical personnel do not need to be notified of all delays of examination but should be notified in those situations where the delay could compromise patient care.
4.5 | Evaluation Tool and Follow Up Is there a tool for regularly evaluating client satisfaction, staff suggestions and is the feedback received effectively utilized to improve services? |  | 2 |  | 0 |  | ISO15189:2012 Clause 4.1.1.4(m); 4.8; 4.14.3; 4.14.4 The laboratory should measure the satisfaction of clients, clinicians and patients regarding its services, either on an ongoing basis or through episodic solicitations.
| Totals |  | 10 |  | 0
Yes
Partial
No
n/a

## SLIPTA_S5
Section 5 - Equipment
Question no. | Question |  | Available Score | Answer | Score | Comments | Standard
5.1 | Adherence to Proper Equipment Protocol |  | 2 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 5.3.1.2 Equipment should be properly placed as specified in user manual away from the following but not limited to water, direct sunlight, vibrations, in traffic and with more than 75% of the base of the equipment sitting on the bench top to avoid tip-over.
|  | Is equipment installed and placed as specified in the operator’s manuals and uniquely labeled or marked? | [5 empty] | 100
|  | Answer to AMR question G5.2 | [5 empty] | 100 | 200
5.2 | Are equipment operated by trained, competent and authorized personnel? |  | 2 |  | 0 |  | ISO15189:2012 Clause 5.3.1.3 Note: The staff must be trained, and deemed competent to operate equipment.
5.3 | Equipment and Method Validation/Verification and Documentation Are newly introduced equipment and methods validated and verified on-site and are validation and verification reports available? |  | 5 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 5.3.1.2; 5.5.1 Newly introduced methods or equipment must be verified onsite to ensure that their introduction yields performance equal to or better than the previous method or equipment. Manufacturers’ validation may be used. Back up equipment must also be included in verification procedures.
|  | a) Are specific verification/validation protocols in place for each equipment and examination procedure? | [5 empty] | 100
|  | b) Is validation performed for all methods laboratory designed or developed methods by the laboratory, standard methods used outside their intended scope, and validated methods that are subsequently modified? | [5 empty] | 100
|  | c) Has validation information been obtained from the manufacturer/method developer as part of the verification? | [5 empty] | 100
|  | d) Have performance characteristics been appropriately selected and evaluated as per intended use? | [5 empty] | 100
|  | e) Were the verification/validation studies appropriate and adequate? | [5 empty] | 100
|  | f) Was the analysis of data appropriate for the selected performance characteristics? | [5 empty] | 100
|  | g) Have the verification/validation results/reports been reviewed and approved by an authorised person? | [5 empty] | 100
|  | Answer to AMR question G5.1 | [5 empty] | 100 | 800
5.4 | Mesurement uncertainty of measured quantity tests Does the laboratory have documented estimates of measurement of uncertainty (UM)? |  | 2 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 5.5.1.4 Note: Measurement of uncertainty should be calculated at different clinical decision levels. Cumulative IQC (minimum 6 months data) may be used to calculate measurement of uncertainty.
|  | a) Has the laboratory calculated the measurement uncertainty for each quantitative measurement procedure? | [5 empty] | 100
|  | b) Has the laboratory defined the performance requirements (factors that affect the UM) for the measurement uncertainty of each measurement procedure and regularly review estimates of measurement uncertainty? | [5 empty] | 100
|  | c) Does the lab make its calculated measurement of uncertainty available to its users upon request? | [5 empty] | 100 | 300
5.5 | Equipment Maintenance Records Is current equipment inventory data available on all equipment in the laboratory? |  | 2 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 4.13; 5.3.1.7 Records shall be maintained for each item of equipment used in the performance of examinations. Such equipment list must include major analysers as well as ancillary equipment like centrifuges, water baths, rotators, fridges, pipettes, timers, printers, computers.
|  | a) Name of equipment | [5 empty] | 100
|  | b) Manufacturer’s or authorized supplier contact details | [5 empty] | 100
|  | c) Condition received (new, used, reconditioned) | [5 empty] | 100
|  | d) Serial number | [5 empty] | 100
|  | e) Date of receiving | [5 empty] | 100
|  | f) Where equipment is obsolete, date when put “out of service” | [5 empty] | 100
|  | g) Date of entry into service after validation / verification) | [5 empty] | 100
|  | h) Location | [5 empty] | 100 | 800
5.6 | Equipment Maintenance Records Is relevant equipment service information readily available in the laboratory? |  | 2 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 4.13; 5.3.1.5; 5.3.1.7 Maintenance records must be maintained for each item of equipment used in the performance of examinations. These records shall be maintained and shall be readily available for the lifespan of the equipment or for any time period required by national, regional and local authorities.
|  | a) Service contract information or service schedule that has been adhered to | [5 empty] | 100
|  | b) Contact details for service provider | [5 empty] | 100
|  | c) Decontamination records before service , repair or decommissioning | [5 empty] | 100
|  | d) Engineer or service provider preventive maintenance records | [5 empty] | 100
|  | e) Last date of service | [5 empty] | 100
|  | f) Next date of service | [5 empty] | 100 | 600
5.7 | Defective Equipment Waiting for Repair Is defective equipment waiting for repair not used and clearly labelled? |  | 2 |  | 0 |  | ISO15189:2012 Clause 4.13; 5.3.1.5 Note: label should include the date of malfunction and “not in use” and a signature of approval.
5.8 | Obsolete Equipment Procedures Is non-functioning equipment appropriately labelled and removed from the laboratory or path of workflow following the equipment management policies and procedures? |  | 2 |  | 0 |  | ISO15189:2012 Clause 4.13; 5.3.1.5 Note: Label should include the date made obsolete and “obsolete” and a signature of approval.
5.9 | Equipment calibration and Metrological traceability Protocol |  | 2 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 5.3.1.4 Note: Documentation of calibration traceability to a higher order reference material or reference procedure may be provided by an examination system manufacturer. Such documentation is acceptable as long as the manufacturer's examination system and calibration procedures are used without modification.
|  | a) Is routine calibration of laboratory equipment (including pipettes, centrifuges, balances and thermometers) scheduled, as indicated on the equipment, and verified? | [5 empty] | 100
|  | b) Is the calibration traceable (e.g. use of reference materials and equipment like certified thermometers, tachometer? | [5 empty] | 100
|  | c) Is there evidence of review of calibration certificates/results by the laboratory before acceptance back into use? | [5 empty] | 100
|  | d) Is certified reference materials, examination and calibration by another procedure, use of mutual consent standards or methods used for in house calibrations? | [5 empty] | 100
|  | Answer to AMR question U8.12 | [5 empty] | 100
|  | Answer to AMR question U8.13 | [5 empty] | 100 | 600
5.10 | Equipment Preventive Maintenance Is routine user preventive maintenance performed on all equipment and recorded according to manufacturer’s minimum requirements? |  | 2 |  | 0 |  | ISO15189:2012 Clause 4.13; 5.3.1.5 Note: Preventative maintenance by operators must be done on all equipment used in examinations including centrifuges, autoclaves, microscopes, and safety cabinets.
5.11 | Equipment Service Maintenance |  | 2 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 4.13; 5.3.1.5 Note: All equipment must be serviced at specified intervals by a qualified service engineer either through service contracts or otherwise. Service schedule must at minimum meet manufacturer’s requirements.
|  | Is equipment routinely serviced according to schedule as per the minimum manufacturer recommendations by qualified and competent personnel and is this information documented in appropriate logs? | [5 empty] | 100
|  | Answer to AMR question G5.3 | [5 empty] | 100 | 200
5.12 | Equipment Malfunction - Response and Documentation Is equipment malfunction resolved by corrective action(s) and the associated root cause analysis? |  | 2 |  | 0 |  | ISO15189:2012 Clause 4.9; 4.10, 4.13; 5.3.1.5 Note: All equipment malfunctions must be investigated and documented as per the non-conforming procedure. In the event that the user cannot resolve the problem, a repair order must be initiated.
5.13 | Equipment Repair Monitoring and Documentation - Are repair orders monitored to determine if the service is completed? - Does the laboratory verify and document that it is in proper working order before being put back into service? |  | 2 |  | 0 |  | ISO15189:2012 Clause 4.13; 5.3.1.5; 5.6 Note: After a repair all levels of QC must or other performance checks must be processed to verify that the equipment is in proper working condition. Copies of the QC or performance checks results should be attached to the repair records as evidence.
5.14 | Equipment Failure - Contingency Plan Is there a functional back-up system that prevents interruption of lab services? |  | 2 |  | 0 |  | ISO15189:2012 Clause 4.1.1.4 (n); 5.3.1 Note: Interruption to services is considered when a laboratory cannot release results to their users. Testing services should not be subject to interruption due to equipment malfunctions. Contingency plans must be in place, in the event of equipment failure, for the completion of testing. In the event of a testing disruption, planning may include the use of a back-up instrument, the use of a different testing method, the referral of samples to another laboratory.
5.15 | Manufacturer's Operator Manual Are the equipment manufacturer’s operator manuals readily available to testing staff, and where possible, available in the language understood by staff? |  | 2 |  | 0 |  | ISO15189:2012 Clause 5.3.1.3 Operator manuals must be readily available for reference by testing staff.
5.16 | Laboratory Testing Services Has the laboratory provided uninterrupted testing services, with no disruptions due to equipment failure in the last year (or since the last audit)? |  | 2 |  | 0 |  | ISO15189:2012 Clause 4.1.1.4(a);(n); 4.1.2.1(i) Note: Interruption to services is considered when a laboratory cannot release results to their users. Testing services should not be subject to interruption due to equipment malfunctions. Contingency plans must be in place, in the event of equipment failure, for the completion of testing. In the event of a testing disruption, planning may include the use of a back-up instrument, the use of a different testing method, the referral of samples to another laboratory
| Totals |  | 35 |  | 0
Yes
Partial
No
N/A

## SLIPTA_S6
Section 6 - Evaluation and Audits
Question no. | Question |  | Available Score | Answer | Score | Comments | Standard
6.1 | Internal audits Are internal audits conducted at intervals as defined in the quality manual and do these audits address areas important to patient care? |  | 5 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 4.13; 4.14.5 Note: The cycle for internal auditing should normally be completed in one year. The laboratory must conduct internal audits at planned intervals to determine whether all activities in the quality management system, including pre-examination, examination, and post-examination.
|  | a) Is there an audit plan/schedule that ensures all activities of the QMS are audited? | [5 empty] | 100
|  | b) Are audits being carried with minimal conflict of interest e.g. where possible, carried out by persons who are not involved in lab activities in the section being audited? | [5 empty] | 100
|  | c) Are the personnel conducting the internal audits trained and competent in auditing? | [5 empty] | 100
|  | d) Is root cause analysis performed for nonconformities/noted deficiencies? | [5 empty] | 100
|  | e) Are internal audit findings documented and presented to the laboratory management and relevant staff for review? | [5 empty] | 100
|  | Answer to AMR question G6.1 | [5 empty] | 100 | 600
6.2 | Audit Recommendations and Action Plan & Follow Up |  | 5 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 4.10; 4.13; 4.14.5 Note: For actions that are not implemented as per the due dates there should be a motivation and an approval of extension.
|  | a) Are internal audits reports generated? | [5 empty] | 100
|  | b) Are recommendations for corrective/preventive actions made based on audit findings? | [5 empty] | 100
|  | c) Is an action plan developed with clear timelines, assigned personnel & documented follow-up within the timeframe defined by the laboratory? | [5 empty] | 100 | 300
6.3 | Risk Management Are assessment of potential pitfalls performed for all laboratory processes including pre examination, examination and post examination? |  | 5 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 4.13; 4.14.6 The Laboratory shall assess all steps in for all its processes (pre-analytical, analytical and post analytical) for areas of potential pitfalls e.g. pre-analytical step of sample collection, potential pitfalls could be; wrong sample collected, sample collected in wrong container, sample collected at wrong time. Post analytical could be; result sent to wrong patient, results sent outside of TAT. The Lab must assess all steps, list potential pitfalls and document action taken to prevent these from occurring. Note: Risks should be graded and acted upon as per their grading.
|  | a) Documented assessment of potential pitfalls for all processes | [5 empty] | 100
|  | b) Documented actions taken to reduce or eliminate identified potential pitfalls | [5 empty] | 100 | 200
| Totals |  | 15 |  | 0
Yes
Partial
No
N/A

## SLIPTA_S7
Section 7 - Purchasing & Inventory
Question no. | Question |  | Available Score | Answer | Score | Comments | Standard
7.1 | Inventory and Budgeting System Is there a system for accurately forecasting needs for supplies and reagents? |  | 2 |  | 0 |  | ISO15189:2012 Clause 4.1.2.1(i); 5.3.2.1; 5.3.2.4 The laboratory must have a systematic way of determining its supply and testing needs through inventory control and budgeting systems that take into consideration past patterns, present trends and future plans.
7.2 | Specifications for supplies and consumables |  | 2 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 4.6 Note: Specification could be in the form of catalogue number; item number, manufacturer name etc.
|  | Does the laboratory provide specification for their supplies and consumables that are required when placing a requisition? | [5 empty] | 100
|  | Answer to AMR question G7.1 | [5 empty] | 100 | 200
7.3 | Service Supplier Performance Review Does the lab monitor the performance of the suppliers to ensure that the stated criteria are met? |  | 2 |  | 0 |  | ISO15189:2012 Clause 4.6 All suppliers of services used by the laboratory must be reviewed for their performance. Those that perform well must be identified and listed as approved suppliers. Results of these reviews must be documented.
7.4 | Inventory Control Does the lab maintain records for each reagent and consumable that contributes to the performance of examinations. These records shall include but not be limited to the following: |  | 2 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 4.13; 5.3.2.7; 5.3.2.4 Note: All incoming orders should be inspected for condition and completeness of the original requests, receipted and documented appropriately; the date received in the laboratory and the expiry date for the product should be clearly indicated.
|  | a) Identity of the reagent or consumable | [5 empty] | 100
|  | b) Batch code or lot number | [5 empty] | 100
|  | c) Manufacturer or supplier name and contact information | [5 empty] | 100
|  | d) Date of receiving, the expiry date, date of entering into service and, where applicable, the date the material was taken out of service | [5 empty] | 100
|  | e) Manufacturer's instruction/package insert | [5 empty] | 100
|  | f) Records of inspection of reagents and consumables when received (e.g. acceptable or damaged) | [5 empty] | 100 | 600
7.5 | Budgetary Projections Are budgetary projections based on personnel, test, facility and equipment needs as well as quality assurance procedures and materials? |  | 2 |  | 0 |  | ISO15189:2012 Clause 4.1.1.4(a) Note: Budgetary projections will ensure that there are no disruptions to services provided
7.6 | Management Review of Supply Requests Does management review/approve the finalized supply requests? |  | 2 |  | 0 |  | ISO15189:2012 Clause 5.3.2.3; 5.3.2.7 Note: Due to the fact that labs have different purchasing approval systems, there should be a system in place that the lab reviews final approval of their original request.
7.7 | Laboratory Inventory System |  | 2 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 5.3.2 Note: The laboratory inventory system should reliably inform staff of the minimum amount of stock to be kept in order to avoid interruption of service due to stock-outs and the maximum amount to be kept by the laboratory to prevent expiry of reagents.
|  | a) Are inventory records complete and accurate, with minimum and maximum stock levels denoted and monitored? | [5 empty] | 100
|  | b) Is the consumption rate of all reagents and consumables monitored? | [5 empty] | 100
|  | c) Are stock counts routinely performed? | [5 empty] | 100 | 300
7.8 | Storage Area Are storage areas set up and monitored appropriately? |  | 2 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 5.3.2.2 Note: Storage of supplies and consumables must be as per the manufacturer’s specifications.
|  | a) Is the storage area well-organized and free of clutter? | [5 empty] | 100
|  | b) Are there designated places for all inventory items for easy access? | [5 empty] | 100
|  | c) Is adequate cold storage available? | [5 empty] | 100
|  | d) Are storage areas monitored as per prescribed storage conditions? | [5 empty] | 100
|  | e) Is the ambient temperature monitored routinely? | [5 empty] | 100
|  | f) Is storage in direct sunlight avoided? | [5 empty] | 100
|  | g) Is the storage area adequately ventilated? | [5 empty] | 100
|  | h) Is the storage area clean and free of dust and pests? | [5 empty] | 100
|  | i) Are storage areas access-controlled? | [5 empty] | 100
|  | Answer to AMR question G7.2 | [5 empty] | 100 | 1000
7.9 | Inventory Organization and Watage Minimization Is First-Expiry-First-Out (FEFO) practised? |  | 2 |  | 0 |  | ISO15189:2012 Clause 5.3.2.2 and USAID Deliver Project, Logistics Handbook, Task Order 1 To minimize wastage from product expiry, inventory should be organized in line with the First-Expiry-First-Out (FEFO) principle. Place products that will expire first in front of products with a later expiry date and issue stock accordingly to ensure products in use are not past their expiry date. Remember that the order in which products are received is not necessarily the order in which they will expire.
7.10 | Product Expiration |  | 2 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 5.3.2.3 All reagents and test kits in use, as well as those in stock, should be within the manufacturer-assigned expiry dates. Expired stock should not be entered into use and should be documented before disposal.
|  | Are all reagents/test kits in use (and in stock) currently within the manufacturer-assigned expiration dates or within stability? | [5 empty] | 100
|  | Answer to AMR question U7.1 | [5 empty] | 100
|  | Answer to AMR question F7.1 | [5 empty] | 100
|  | Answer to AMR question B7.1 | [5 empty] | 100
|  | Answer to AMR question N7.1 | [5 empty] | 100
|  | Answer to AMR question P7.1 | [5 empty] | 100
|  | Answer to AMR question W7.1 | [5 empty] | 100 | 700
7.12 | Disposal of Expired Products Are expired products labeled and disposed of properly? |  | 2 |  | 0 |  | ISO15189:2012 Clause 5.3.2.7 Expired products should be disposed of properly and records maintained. If safe disposal is not available at the laboratory, the manufacturer/supplier should take back the expired stock at the time of their next delivery.
7.13 | Laboratory Testing Services Has the laboratory provided uninterrupted testing services, with no disruptions due to stock-outs in the last year or since last audit? |  | 2 |  | 0 |  | ISO15189:2012 Clause 4.1.1.4(a);(n); 4.1.2.1(i); 5.5 Note: Interruption to services is considered when a laboratory cannot release results to their users. Testing services should not be subject to interruption due to stock-outs. Laboratories should pursue all options for borrowing stock from another laboratory or referring samples to another testing facility while the stock-out is being addressed.
| Totals |  | 24 |  | 0
Yes
Partial
No
N/A

## SLIPTA_S8
Section 8 - Process Control
Question no. | Question |  | Available Score | Answer | Score | Comments | Standard
8.1 | Information for patients and users Are guidelines for patient identification, specimen collection (including client safety), labeling and transport readily available to persons responsible for primary sample collection? |  | 2 |  | 0 |  | ISO15189:2012 Clause 5.4.1 Note: The laboratory shall have documented procedures and information for pre-examination activities to ensure the validity of the results of examinations and must make these available to those who collect samples.
8.2 | Does the laboratory adequately collect information needed for examination performance? |  | 3 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 4.4; 5.4.3 Note: Each request accepted by the laboratory for examination(s) shall be considered an agreement. The request may be in the form of a hard copy or electronically.
|  | a) Are all test requests accompanied by an acceptable and approved test requisition form (and a transmittal sheet/checklist/manifest where applicable)? | [5 empty] | 100
|  | b) Does the request form have patient ID including gender, date of birth, location of patient and unique identifier? | [5 empty] | 100
|  | c) Name, signature or initials of authorized requester | [5 empty] | 100
|  | d) Type of sample and examination requested | [5 empty] | 100
|  | e) Clinically relevant information | [5 empty] | 100
|  | f) Date of sample collection (And time of collection where relevant – where time has an impact on the result) | [5 empty] | 100
|  | g) Date and time of sample receipt | [5 empty] | 100
|  | h) Written consent for invasive procedures with increased risk of complications | [5 empty] | 100
|  | Answer to AMR question G8.1 | [5 empty] | 100
|  | Answer to AMR question U8.1 | [5 empty] | 100 | 1000
8.3 | Are adequate sample receiving procedures in place? |  | 2 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 4.4; 5.4.6 Note: The review of service agreements occurs on sample reception. All portions of the primary sample must be unequivocally traceable to the original primary sample.
|  | a) Patient Unique Identifier? | [5 empty] | 100
|  | b) Are received specimens evaluated according to acceptance/rejection criteria? | [5 empty] | 100
|  | c) Are specimens logged appropriately upon receipt in the laboratory (including date, time, and name of receiving officer)? | [5 empty] | 100
|  | d) Are procedures in place to process “urgent” specimens and verbal requests? | [5 empty] | 100
|  | e) When samples are split, can the portions be traced back to the primary sample? | [5 empty] | 100
|  | f) If not a 24-hour lab, is there a documented method for handling specimens received after hours? | [5 empty] | 100
|  | g) Are specimens delivered to the correct workstations in a timely manner? | [5 empty] | 100 | 700
8.4 | Pre-examination Handling, Preparation and Storage Where testing does not occur immediately upon arrival in the laboratory, are specimens stored appropriately prior to testing? |  | 2 |  | 0 |  | ISO15189:2012 Clause 5.4.7 Note: Specimens should be stored under the appropriate conditions to maintain the stability of the specimen.
8.5 | Sample Transportation |  | 2 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 5.4.4.3; 5.4.5 Note: All samples should be transported to the laboratory in a manner that is safe to the patients, the public and the environment. The laboratory must ensure that the samples were received within a temperature interval specified for sample collection.
|  | Are specimens, either received or referred, packaged appropriately according to local and/or international regulations and transported within acceptable timeframes and temperature intervals? | [5 empty] | 100
|  | Answer to AMR question U8.2 | [5 empty] | 100
|  | Answer to AMR question U8.6 | [5 empty] | 100
|  | Answer to AMR question F8.4 | [5 empty] | 100
|  | Answer to AMR question B8.4 | [5 empty] | 100 | 500
8.6 | Does the laboratory select and evaluate referral Labs and Consultants? |  | 2 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 4.13; 4.5 Note: The laboratory must have system in place to ensure that the referral laboratories are competent to perform the services required. Evaluations in the form of checking their accreditation status, using a questionnaire, performing audits, use of blinded samples etc.
|  | a) Are there documented reviews and evaluations of referral laboratories and consultants as defined by the laboratory? | [5 empty] | 100
|  | b) Is there a register of referral Laboratories and consultants? | [5 empty] | 100
|  | c) Are referred specimens tracked properly using a logbook, tracking form or electronically? | [5 empty] | 100 | 300
8.7 | Documentation of Examination Procedures |  | 2 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 5.5.3 Note: examination procedures are for the laboratory staff to use therefore it should be in the language that is commonly understood by the staff; the lab may translate the documents into other languages which must be document controlled.
|  | Are examination procedures documented in a language commonly understood by all staff and available in appropriate locations? | [5 empty] | 100
|  | Summary urine questions |  | Answer all sub questions |  |  |  |  | 100
|  | Answer to AMR question U8.7 | [5 empty] | 100
|  | Answer to AMR question U8.8 | [5 empty] | 100
|  | Answer to AMR question U8.9 | [5 empty] | 100
|  | Answer to AMR question U8.10 | [5 empty] | 100
|  | Answer to AMR question U8.11 | [5 empty] | 100
|  | Answer to AMR question U8.14 | [5 empty] | 100
|  | Answer to AMR question U8.15 | [5 empty] | 100
|  | Answer to AMR question U8.16 | [5 empty] | 100
|  | Answer to AMR question U8.17 | [5 empty] | 100
|  | Answer to AMR question U8.18 | [5 empty] | 100
|  | Answer to AMR question U8.19 | [5 empty] | 100
|  | Answer to AMR question U8.20 | [5 empty] | 100
|  | Answer to AMR question U8.21 | [5 empty] | 100
|  | Answer to AMR question U8.22 | [5 empty] | 100 | 1400
|  | Summary feces questions |  | Answer all sub questions |  |  |  |  | 100
|  | Answer to AMR question F8.5 | [5 empty] | 100
|  | Answer to AMR question F8.6 | [5 empty] | 100
|  | Answer to AMR question F8.7 | [5 empty] | 100
|  | Answer to AMR question F8.8 | [5 empty] | 100
|  | Answer to AMR question F8.9 | [5 empty] | 100
|  | Answer to AMR question F8.10 | [5 empty] | 100
|  | Answer to AMR question F8.11 | [5 empty] | 100
|  | Answer to AMR question F8.12 | [5 empty] | 100
|  | Answer to AMR question F8.13 | [5 empty] | 100
|  | Answer to AMR question F8.14 | [5 empty] | 100
|  | Answer to AMR question F8.15 | [5 empty] | 100
|  | Answer to AMR question F8.16 | [5 empty] | 100
|  | Answer to AMR question F8.17 | [5 empty] | 100
|  | Answer to AMR question F8.18 | [5 empty] | 100 | 1400
|  | Summary blood questions |  | Answer all sub questions |  |  |  |  | 100
|  | Answer to AMR question B8.5 | [5 empty] | 100
|  | Answer to AMR question B8.6 | [5 empty] | 100
|  | Answer to AMR question B8.7 | [5 empty] | 100
|  | Answer to AMR question B8.8 | [5 empty] | 100
|  | Answer to AMR question B8.9 | [5 empty] | 100
|  | Answer to AMR question B8.10 | [5 empty] | 100
|  | Answer to AMR question B8.11 | [5 empty] | 100
|  | Answer to AMR question B8.12 | [5 empty] | 100
|  | Answer to AMR question B8.13 | [5 empty] | 100
|  | Answer to AMR question B8.14 | [5 empty] | 100
|  | Answer to AMR question B8.15 | [5 empty] | 100
|  | Answer to AMR question B8.16 | [5 empty] | 100
|  | Answer to AMR question B8.17 | [5 empty] | 100
|  | Answer to AMR question B8.18 | [5 empty] | 100
|  | Answer to AMR question B8.19 | [5 empty] | 100 | 1500
|  | Summary genital questions |  | Answer all sub questions |  |  |  |  | 100
|  | Answer to AMR question N8.10 | [5 empty] | 100
|  | Answer to AMR question N8.11 | [5 empty] | 100
|  | Answer to AMR question N8.12 | [5 empty] | 100
|  | Answer to AMR question N8.13 | [5 empty] | 100
|  | Answer to AMR question N8.14 | [5 empty] | 100
|  | Answer to AMR question N8.15 | [5 empty] | 100
|  | Answer to AMR question N8.16 | [5 empty] | 100 | 700
|  | Summary pulmonary questions |  | Answer all sub questions |  |  |  |  | 100
|  | Answer to AMR question P8.4 | [5 empty] | 100
|  | Answer to AMR question P8.5 | [5 empty] | 100
|  | Answer to AMR question P8.6 | [5 empty] | 100
|  | Answer to AMR question P8.7 | [5 empty] | 100
|  | Answer to AMR question P8.8 | [5 empty] | 100
|  | Answer to AMR question P8.9 | [5 empty] | 100
|  | Answer to AMR question P8.10 | [5 empty] | 100
|  | Answer to AMR question P8.11 | [5 empty] | 100
|  | Answer to AMR question P8.12 | [5 empty] | 100
|  | Answer to AMR question P8.13 | [5 empty] | 100
|  | Answer to AMR question P8.14 | [5 empty] | 100
|  | Answer to AMR question P8.15 | [5 empty] | 100
|  | Answer to AMR question P8.16 | [5 empty] | 100
|  | Answer to AMR question P8.17 | [5 empty] | 100
|  | Answer to AMR question P8.18 | [5 empty] | 100
|  | Answer to AMR question P8.19 | [5 empty] | 100
|  | Answer to AMR question P8.20 | [5 empty] | 100
|  | Answer to AMR question P8.21 | [5 empty] | 100 | 1800
|  | Summary wound questions |  | Answer all sub questions |  |  |  |  | 100
|  | Answer to AMR question W8.4 | [5 empty] | 100
|  | Answer to AMR question W8.5 | [5 empty] | 100
|  | Answer to AMR question W8.6 | [5 empty] | 100
|  | Answer to AMR question W8.7 | [5 empty] | 100
|  | Answer to AMR question W8.8 | [5 empty] | 100
|  | Answer to AMR question W8.9 | [5 empty] | 100
|  | Answer to AMR question W8.10 | [5 empty] | 100
|  | Answer to AMR question W8.11 | [5 empty] | 100
|  | Answer to AMR question W8.12 | [5 empty] | 100
|  | Answer to AMR question W8.13 | [5 empty] | 100
|  | Answer to AMR question W8.14 | [5 empty] | 100 | 1200
|  | Answer to AMR question W8.15 | [5 empty] | 100 | 700
8.8 | Reagents Acceptance Testing |  | 2 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 5.3.2.3 Note: This may be accomplished by a comparison study or examining quality control samples and verifying that results are acceptable.
|  | Is each new reagent preparation, lot number, new shipment of reagents, or consumable verified before use and documented? | [5 empty] | 100
|  | Answer to AMR question G8.2 | [5 empty] | 100
|  | Answer to AMR question U8.3 | [5 empty] | 100
|  | Answer to AMR question U8.4 | [5 empty] | 100
|  | Answer to AMR question F8.1 | [5 empty] | 100
|  | Answer to AMR question F8.2 | [5 empty] | 100
|  | Answer to AMR question B8.1 | [5 empty] | 100
|  | Answer to AMR question B8.2 | [5 empty] | 100
|  | Answer to AMR question N8.7 | [5 empty] | 100
|  | Answer to AMR question N8.8 | [5 empty] | 100
|  | Answer to AMR question P8.1 | [5 empty] | 100
|  | Answer to AMR question P8.2 | [5 empty] | 100
|  | Answer to AMR question W8.1 | [5 empty] | 100
|  | Answer to AMR question W8.2 | [5 empty] | 100 | 1400
8.9 | Quality Control |  | 3 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 5.6.2 The laboratory shall design internal quality control systems that verify the attainment of the intended quality of results.
|  | Is internal quality control performed, documented and verified for all tests/procedures before releasing patient results? | [5 empty] | 100
|  | Answer to AMR question G8.6 | [5 empty] | 100 | 200
8.10 | Quality Control Data Are QC results monitored and reviewed (biases, shifts, trends and Levy-Jennings charts)? |  | 3 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 5.6.2.3 Note: The lab must document and implement a system it would use to evaluate patient results since the last successful quality control; the evaluation could be done by re-examining selected samples of various batches, re-examining samples as per the stability of the Quality Control etc.
|  | a) Is there documentation of timely corrective action taken when quality control results exceed the acceptable range or reviews identify nonconformities? | [5 empty] | 100
|  | Answer to AMR question G8.3 | [5 empty] | 100
|  | Answer to AMR question G8.4 | [5 empty] | 100
|  | Answer to AMR question G8.5 | [5 empty] | 100
|  | b) Does the Lab evaluate the results from the patient samples that were examined after the last successful quality control event | [5 empty] | 100
|  | Answer to AMR question U8.5 | [5 empty] | 100
|  | Answer to AMR question F8.3 | [5 empty] | 100
|  | Answer to AMR question B8.3 | [5 empty] | 100
|  | Answer to AMR question N8.9 | [5 empty] | 100
|  | Answer to AMR question P8.3 | [5 empty] | 100
|  | Answer to AMR question W8.3 | [5 empty] | 100 | 1100
8.11 | Comparability of Examination Results Does the laboratory compare results of the same test performed with different procedures and equipment? |  | 2 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 5.6.4 Note: The lab should document and implement a system to ensure there is comparability of results, this could be done by the use of EQA performance; using blinded samples, parallel testing.
|  | a) Where there is more than one procedure for the same measure, does the laboratory compare results from the different procedures, equipment or methods? | [5 empty] | 100
|  | b) Does the lab discuss, document and act upon (including notifying users) problems or deficiencies from these comparison studies? | [5 empty] | 100 | 200
8.12 | Are environmental conditions checked and reviewed accurately? Are the following environmental conditions checked daily? |  | 2 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 5.4.1 Note: The laboratory shall monitor, control and record environmental conditions, as required by relevant specifications or where they may influence the quality of the results and/or the health of staff.
|  | a) Room temperature | [5 empty] | 100
|  | b) Freezer(s) | [5 empty] | 100
|  | c) Refrigerator(s) | [5 empty] | 100
|  | d) Incubator(s) | [5 empty] | 100
|  | e) Water bath(s) | [5 empty] | 100 | 500
8.13 | Have acceptable ranges been defined for all temperature-dependent equipment with procedures and documentation of action taken in response to out-of-range temperatures? |  | 2 |  | 0 |  | ISO15189:2012 Clause 5.2.2(c) Note: Acceptable ranges should take into consideration manufacturers’ recommendations and requirements.
8.14 | Does the laboratory participate in inter-laboratory comparison program or alternative assessment systems for all tests? |  | 3 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 5.6.3 Note: The laboratory should handle, analyze, review and report results for proficiency testing in a manner similar to regular patient testing. Investigation and correction of problems identified by unacceptable proficiency testing should be documented. Acceptable results showing bias or trends suggest that a problem should also be investigated.
|  | a) Do samples come from providers who are accredited or approved? | [5 empty] | 100
|  | b) Are specimens handled and tested the same way as patient specimens? | [5 empty] | 100
|  | c) Is the performance of the laboratory in the PT program reviewed and discussed with relevant staff? | [5 empty] | 100
|  | d) Is cause analysis performed for unacceptable results? | [5 empty] | 100
|  | e) Is corrective action documented for unacceptable results? | [5 empty] | 100
|  | Answer to AMR question U8.21 | [5 empty] | 100
|  | Answer to AMR question U8.22 | [5 empty] | 100
|  | Answer to AMR question U8.23 | [5 empty] | 100
|  | Answer to AMR question F8.19 | [5 empty] | 100
|  | Answer to AMR question F8.20 | [5 empty] | 100
|  | Answer to AMR question F8.21 | [5 empty] | 100
|  | Answer to AMR question B8.20 | [5 empty] | 100
|  | Answer to AMR question B8.21 | [5 empty] | 100
|  | Answer to AMR question B8.22 | [5 empty] | 100
|  | Answer to AMR question N8.17 | [5 empty] | 100
|  | Answer to AMR question N8.18 | [5 empty] | 100
|  | Answer to AMR question N8.19 | [5 empty] | 100
|  | Answer to AMR question P8.22 | [5 empty] | 100
|  | Answer to AMR question P8.23 | [5 empty] | 100
|  | Answer to AMR question P8.24 | [5 empty] | 100
|  | Answer to AMR question W8.22 | [5 empty] | 100
|  | Answer to AMR question W8.23 | [5 empty] | 100
|  | Answer to AMR question W8.24 | [5 empty] | 100 | 2300
| Totals |  | 32 |  | 0
Yes
Partial
No
N/A

## SLIPTA_S9
Section 9 - Information Management
Question no. | Question |  | Avaliable Score | Answer | Score | Comments | Standard
9.1 | Test Result Reporting System Are test results legible, technically verified by an authorized person and confirmed against patient identity? |  | 2 |  | 0 |  | ISO15189:2012 Clause 5.8.1 Results must be written in ink and written clearly with no mistakes in transcription. Cancellation must follow Good Lab Practices. The persons performing the test must indicate verification of the results. There must be a signature or identification of the person authorizing the release of the report.
9.1 | Testing Personnel Are testing personnel identified on the result report or other records (manual or electronic)? |  | 2 |  | 0 |  | ISO15189:2012 Clause 4.13 ; 5.5.1.1; 5.8.1 Note: The person who performed the procedure must be identified on the report (hard copy or electronic) purposes of traceability.
9.2 | Report Content Does the laboratory report contain at least the following: |  | 3 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 5.8.2; 5.8.3; 5.9.3 Note: When the reporting system cannot capture amendments, changes or alterations, a record of such shall be kept.
|  | a) Test requested | [5 empty] | 100
|  | b) Identification of the laboratory | [5 empty] | 100
|  | c) Identification of all examinations performed by a referral laboratory | [5 empty] | 100
|  | c) Patient identification and location | [5 empty] | 100
|  | d) Name of the requester | [5 empty] | 100
|  | e) Date of primary sample collection (and time, if relevant to patient care) | [5 empty] | 100
|  | f) type of primary sample | [5 empty] | 100
|  | g) Is the result reported in SI units where applicable? | [5 empty] | 100
|  | h) Biological reference intervals where applicable | [5 empty] | 100
|  | i) Is there space for interpretation or comments of results, when applicable? | [5 empty] | 100
|  | j) Identification of the person(s) reviewing and authorizing the report | [5 empty] | 100
|  | k) Date and time of the report | [5 empty] | 100
|  | l) Page number to total number of pages (e.g. "Page 1 of 5", "Page 2 of 5", etc.) | [5 empty] | 100
|  | m) When issuing revised reports, is it clearly identified as a revision and includes reference to the date and patient's identity in the original report and the user made aware of the revision? | [5 empty] | 100
|  | n) Does the revised record show the time and date of the change and the name of the person responsible for the change? | [5 empty] | 100
|  | o) Does the original report entry remain in the record when revisions are made? | [5 empty] | 100 | 1600
9.3 | Analytic System/Method Tracing |  | 2 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 4.13(g) Note: There must be traceability of specimen results to a specific analytical system or method. Proficiency testing specimens would also fall under specimen results.
|  | When more than one instrument is in use for the same test, are test results traceable to the equipment used for testing? | [5 empty] | 100
|  | Answer to AMR question G9.1 | [5 empty] | 100
|  | Answer to AMR question U9.1 | [5 empty] | 100
|  | Answer to AMR question U9.2 | [5 empty] | 100
|  | Answer to AMR question F9.1 | [5 empty] | 100
|  | Answer to AMR question F9.2 | [5 empty] | 100
|  | Answer to AMR question B9.1 | [5 empty] | 100
|  | Answer to AMR question B9.2 | [5 empty] | 100
|  | Answer to AMR question N9.1 | [5 empty] | 100
|  | Answer to AMR question P9.1 | [5 empty] | 100
|  | Answer to AMR question P9.2 | [5 empty] | 100
|  | Answer to AMR question W9.1 | [5 empty] | 100
|  | Answer to AMR question W9.2 | [5 empty] | 100 | 1300
9.4 | Archived data Labeling and Storage Are archived results (paper or data-storage media) properly labeled and stored in a secure location accessible only to authorized personnel? |  | 2 |  | 0 |  | ISO15189:2012 Clause 4.13; 5.10.3 All patient data, paper, tapes, disks should be properly labeled and stored securely in places accessible only to authorized personnel. ISO 15189: 5.8.3 Annex B 6.4.
9.5 | Authorities and Responsibilities Has the laboratory defined and implemented authorities and responsibilities for the management and use of the laboratory information system– paper based and electronic, including maintenance and modifications that may affect patient care? |  | 2 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 5.9; 5.10.2; 5.10.3 Note:"information systems" includes the management of data and information contained in both computer and non-computerized systems. Some of the requirements may be more applicable to computer systems than to non-computerized systems. Computerized systems can include those integral to the functioning of laboratory equipment and standalone systems using generic software, such as word processing, spreadsheet and database applications that generate, collate, report and archive patient information and reports.
Is the following in place and implemented?
|  | a) Controlled access to patient data and information | [5 empty] | 100
|  | b) Controlled access to enter patient data and examination results | [5 empty] | 100
|  | c) Controlled access to changing patient data or examination results | [5 empty] | 100
|  | d) Controlled access to the release of examination results and reports | [5 empty] | 100
|  | e) Verification that results that have been transmitted electronically or reproduced external to the laboratory (computers, fax machines, email and websites and personal web devices) are correct. | [5 empty] | 100 | 500
9.6 | Information Management System Does the laboratory have evidence of how the Laboratory Information Management System (LIMS) was selected? |  | 2 |  | 0 |  | ISO15189:2012 Clause 5.3.1.1 Note: The laboratory must have a documented procedure and records for the selection, purchasing and management of equipment.
9.7 | Test Result Are test results validated, interpreted and released by appropriately authorized personnel? |  | 2 |  | 0 |  | ISO15189:2012 Clause 5.1; 5.8; 5.10.3; 5.9.1 Note: There must be a signature or identification of the person authorizing the release of the report.
9.8 | Verification of Electronic Laboratory Information System |  | 2 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 4.13; 5.10.3 Note: The lab must perform verification of system after upgrades and to ensure previously stored patient results have not been affected.
If not applicaple select N/A for all sub questions
|  | a) Has the system been verified before implementation that include the verification reports to check functioning and inter-phasing by the laboratory? | [5 empty] | 100
|  | b) Are records of the validation by the supplier available and approved for use? | [5 empty] | 100
|  | c) Are reports of ongoing system checks available for correct transmissions, calculations and storage of results and records. | [5 empty] | 100 | 300
9.9 | Is the Laboratory Information System properly maintained to ensure continued functioning: |  | 2 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 5.10.3 Note: If the LIS is maintained offsite, records of maintenance must be readily available .The lab should include the LIS as part of their internal audit.
If not applicaple select N/A for all sub questions
|  | a) Documented regular service by authorized and trained personnel | [5 empty] | 100
|  | b) Documented system failures with documented appropriate root cause analysis, corrective actions and preventive actions | [5 empty] | 100
|  | c) System operated in an environment recommended by the supplier for optimal functioning | [5 empty] | 100 | 300
| Totals |  | 21 |  | 0
Yes
Partial
No
N/A

## SLIPTA_S10
Section 10 - Identification of Non Conformities, Corrective and Preventive Actions
Question no. | Question |  | Avaliable Score | Answer | Score | Comments | Standard
10.1 | Are all identified nonconforming activities/ work identified and documented adequately |  | 5 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 4.9 Note: nonconformities should be identified and managed in any aspect of the quality management system, including pre-examination, examination or post-examination processes. Nonconforming examinations or activities occur in many different areas and can be identified in many different ways, including clinician complaints, internal quality control indications, and instrument calibrations, checking of consumable materials, inter-laboratory comparisons, staff comments, reporting and certificate checking, laboratory management reviews, and internal and external audits.
|  | a) Indicating details of what happened, when, reporting person | [5 empty] | 100
|  | b) Immediate actions being taken | [5 empty] | 100
|  | c) Determination of the extent of the non- conformity? | [5 empty] | 100
|  | d) Are examinations halted and results withheld or recalled where the non-conformity compromises patient results? | [5 empty] | 100
|  | e) Informing the requester where the non-conformity has an effect on the management of the patient | [5 empty] | 100
|  | f) Authorization of resumption of testing documented (where testing has been halted) | [5 empty] | 100
|  | Answer to AMR question G10.1 | [5 empty] | 100 | 700
10.2 | Root cause analysis |  | 3 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 4.10(b) Note: Root cause analysis is a process of identifying and removing the underlying factor of the non-conformance.
|  | Is documented root cause analysis performed on nonconformities before corrective actions are implemented? | [5 empty] | 100
|  | Answer to AMR question G10.2 | [5 empty] | 100 | 200
10.3 | Is corrective action performed and documented for nonconformities? |  | 3 |  | 0 |  | ISO15189:2012 Clause 4.10; 4.13; 4.14.5 Note: Documenting corrective action allows the lab to review its effectiveness and to perform trend analysis for continual improvement.
10.4 | Are implemented corrective actions monitored and reviewed for their effectiveness before closure/clearance? |  | 3 |  | 0 |  | ISO15189:2012 Clause 4.10(f) Note: Implemented corrective action does not imply effectiveness; therefore the lab has to monitor to ensure that the NC has not recurred.
10.5 | Preventive Actions Are documented preventive actions implemented and monitored for their effectiveness? |  | 5 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 4.11; 4.12; Note: Preventive action should be an ongoing process involving analysis of laboratory data, including trend and risk analyses and external quality assessment (proficiency testing).
|  | a) Reviewing of laboratory data and information to determine potential non conformities | [5 empty] | 100
|  | b) Determining root causes for potential non conformities | [5 empty] | 100
|  | c) Implementing and documenting preventive actions | [5 empty] | 100
|  | d) Reviewing and documenting effectiveness of preventive actions | [5 empty] | 100 | 400
| Totals |  | 19 |  | 0
Yes
Partial
No

## SLIPTA_S11
Section 11 - Occurrence/Incident Management & Process Improvement
Question no. | Question |  | Available Score | Answer | Score | Comments | Standard
11.1 | Are graphical tools (charts and graphs) used to communicate quality findings and identify trends? |  | 2 |  | 0 |  | ISO15189:2012 Clause 4.12; 4.13; 4.14 Use of graphical displays of quality data communicates more effectively than tables of numbers. Examples of graphical tools commonly used for this purpose include Pareto charts, cause-and-effect diagrams, frequency histograms, trend graphs, and flow charts.
11.2 | Quality Management System Improvement Measures Does the laboratory identify and undertake continual quality improvement projects? |  | 2 |  | 0 |  | ISO15189:2012 Clause 4.12; 4.15 Note: The lab should use its management review activities to continually improve its quality management system by comparing its actual performance to its intentions stated in the quality policy and objectives.
11.3 | Communication System on Laboratory Operations Does the laboratory communicate with upper management regularly regarding needs for continual improvement? |  | 2 |  | 0 |  | ISO15189:2012 Clause 4.15.2 (o) Note: The laboratory staff should give input for management meetings.
11.4 | Quality indicators |  | 2 | Answer all sub questions | 0 |  | ISO15189:2012 Clause4.12; 4.14.7 Note: The lab should select QI in line with meeting its objectives from pre-analytic, analytic and post-analytic phases critical to patient outcomes.
|  | Are quality indicators (TAT, rejected specimens, stock-outs, etc.) selected, tracked and reviewed? | [5 empty] | 100
|  | Answer to AMR question G11.1 | [5 empty] | 100
|  | Answer to AMR question G11.2 | [5 empty] | 100
|  | Answer to AMR question U11.1 | [5 empty] | 100
|  | Answer to AMR question F11.1 | [5 empty] | 100
|  | Answer to AMR question B11.1 | [5 empty] | 100
|  | Answer to AMR question N11.1 | [5 empty] | 100
|  | Answer to AMR question P11.1 | [5 empty] | 100
|  | Answer to AMR question W11.1 | [5 empty] | 100 | 900
11.5 | Is the outcome of the review of quality indicators used to improve lab performance? |  | 2 |  | 0 |  | ISO15189:2012 Clause 4.14.7; 4.15.2(f) Note: The lab should review the QI to ensure its continued appropriateness.
11.6 | Are the actions taken checked and monitored to determine the effectiveness of improved quality of lab performance? |  | 2 |  | 0 |  | ISO15189:2012 Clause 4.14.7 Note: the lab should create an action plan to monitor the QI stating the objectives, methodology, interpretation, limits, action plan and duration of measurement for each QI.
| Totals |  | 12 |  | 0
Yes
Partial
No

## SLIPTA_S12
Section 12 - Facilities and Safety
Question no. | Question |  | Available Score | Answer | Score | Comments | Standard
12.1 | Is there documented evidence that the laboratory has evaluated the adequacy of the size and overall layout of the laboratory and organized the space so that workstations are positioned for optimal workflow? |  | 2 |  | 0 |  | ISO15189:2012 Clause 5.2.1 Note: Documentation could be in the form of a floor plan, results from internal audits, etc.
12.2 | Are the patient care and testing areas of the laboratory distinctly separate from one another? |  | 2 |  | 0 |  | ISO15189:2012 Clause 5.2.1 Client service areas (i.e. waiting room, phlebotomy room) should be distinctly separate from the testing areas of the laboratory. Client access should not compromise “clean” areas of the laboratory. For biosafety reasons, microbiology and TB testing should be segregated in a separate room(s) from the general laboratory testing.
12.3 | Is each individual workstation maintained free of clutter and set up for efficient operation? |  | 2 | Answer all sub questions | 0 |  | ISO15190 Clause 6.3.5
Are the following criteria met?
|  | a) Does the equipment placement/layout facilitate optimum workflow? | [5 empty] | 100
|  | b) Are all needed supplies present and easily accessible? | [5 empty] | 100
|  | c) Are the chairs/stools at the workstations appropriate for bench height and the testing operations being performed? | [5 empty] | 100 | 300
12.4 | Is the physical work environment appropriate for testing? |  | 2 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 5.2 Note: The laboratory space should be sufficient to ensure the quality of work, safety of personnel and the ability of staff to carry out their tasks without compromising the quality of the examinations. The laboratory should be clean and well organized, free of clutter, well ventilated, adequately lit and within acceptable temperature ranges.
|  | a) Free of clutter? |  |  |  |  | ISO 15190:2003 Clause 13.0 | 100
|  | b) Adequately ventilated? |  |  |  |  | ISO 15190:2003 Clause 6.3.3 | 100
|  | c) Adequately lit? |  |  |  |  | ISO 15190:2003 Clause 6.3.1 | 100
|  | d) Climate-controlled for optimum equipment function? |  |  |  |  | ISO 15190:2003 Clause 6.3.2 | 100
|  | e) Are filters checked, cleaned and/or replaced at regular intervals, where air-conditioning is installed? | [5 empty] | 100
|  | f) Are wires and cables properly located and protected from traffic? | [5 empty] | 100
|  | g) Is there a functioning back-up power supply (generator)? | [5 empty] | 100
|  | h) Is critical equipment supported by uninterrupted power source (UPS) systems? | [5 empty] | 100
|  | i) Is equipment placed appropriately (away from water hazards, out of traffic areas)? | [5 empty] | 100
|  | j) Are appropriate provisions made for adequate water supply, including deionized or distilled water, if needed? | [5 empty] | 100
|  | k) Is clerical work completed outside the testing area? | [5 empty] | 100
|  | l) Is major safety signage posted and enforced, including NO EATING, SMOKING, DRINKING? | [5 empty] | 100
|  | Answer to AMR question G12.2 | [5 empty] | 100 | 1300
12.5 | Laboratory Access Is the laboratory properly secured from unauthorized access with appropriate signage? |  | 2 |  | 0 |  | ISO15189:2012 Clause 5.2.2 Note: Access control should take into consideration safety, confidentiality, and quality.
12.6 | Laboratory Storage Areas Is laboratory-dedicated cold and room temperature storage free of staff food items, and are patient samples stored separately from reagents and blood products in the laboratory refrigerators and freezers? |  | 2 |  | 0 |  | ISO15189:2012 Clause 5.2; 5.2.4 Note: there should be effective separation to prevent contamination.
12.7 | Cleaning & disinfection Is the work area clean and free of leakage and spills, and are disinfection procedures conducted and documented? |  | 2 |  | 0 |  | ISO15189:2012 Clause 5.2.6 Note: The work area should be cleaned regularly. An appropriate disinfectant should be used. At a minimum, all bench tops and working surfaces should be disinfected at the beginning and end of every shift. All spills should be contained immediately and the work surfaces disinfected.
12.8 | Biosafety Cabinet |  | 2 | Answer all sub questions | 0 |  | ISO 15189:2012 Clause 5.2.1; 5.2.2 Note: A biosafety cabinet should be used to prevent aerosol exposure to contagious specimens or organisms. For proper functioning and full protection, biosafety cabinets require periodic maintenance and should be serviced accordingly. Biosafety cabinet should be recertified according to national protocol or manufacturer requirements.
|  | Where a Biosafety cabinet is required to perform work, is it certified and appropriate? | [5 empty] | 100
|  | Answer to AMR question B12.1 | [5 empty] | 100
|  | Answer to AMR question P12.1 | [5 empty] | 100 | 300
12.9 | Laboratory Safety Manual Is a laboratory safety manual available, accessible and up-to-date? |  | 2 | Answer all sub questions | 0 |  | ISO 15190: 7.4 A safety manual should be readily available in work areas as required reading for all employees. The manual should be specific to the laboratory's needs; it should be reviewed and updated at least annually by laboratory management.
Does the safety manual include guidelines on the following topics?
|  | a) Blood and body fluid precautions | [5 empty] | 100
|  | b) Hazardous waste disposal | [5 empty] | 100
|  | c) Hazardous chemicals/materials | [5 empty] | 100
|  | d) MSDS sheets | [5 empty] | 100
|  | e) Personal protective equipment | [5 empty] | 100
|  | f) Vaccination | [5 empty] | 100
|  | g) Post-exposure prophylaxis | [5 empty] | 100
|  | h) Fire safety | [5 empty] | 100
|  | i) Electrical safety | [5 empty] | 100 | 900
12.10 | Waste disposal Is sufficient waste disposal capacity available and is waste separated into infectious and non-infectious waste, with infectious waste autoclaved/incinerated? |  | 2 |  | 0 |  | ISO 15190: 22 Waste should be separated according to biohazard risk, with infectious and non-infectious waste disposed of in separate containers. Infectious waste should be discarded into containers that do not leak and are clearly marked with a biohazard symbol. Sharp instruments and needles should be discarded in puncture resistant containers. Both infectious waste and sharps containers should be autoclaved before being discarded to decontaminate potentially infectious material. To prevent injury from exposed waste, infectious waste should be incinerated, burnt in a pit or buried.
12.11 | Hazardous chemicals Are hazardous chemicals/materials properly handled? |  | 2 | Answer all sub questions | 0 |  | ISO 15190: 17.1 and 17.3 All hazardous chemicals must be labeled with the chemical’s name and with hazard markings clearly indicated. Flammable chemicals must be stored out of sunlight and below their flashpoint, preferably in a still cabinet in a well-ventilated area. Flammable and corrosive agents should be separated from one another. Distinct care should always be taken to handle hazardous chemicals safety in the workplace.
|  | a) Are hazardous chemicals properly labelled? | [5 empty] | 100
|  | b) Are hazardous chemicals properly stored to ensure safety and prevent theft? | [5 empty] | 100
|  | c) Are hazardous chemicals properly utilized according to MSDS? | [5 empty] | 100
|  | d) Are hazardous chemicals properly disposed according to national guidelines or MSDS? | [5 empty] | 100 | 400
12.12 | Handling of sharps Are ‘sharps’ handled and disposed of properly in ‘sharps’ containers that are appropriately utilized? |  | 2 |  | 0 |  | ISO15189:2012 Clause 5.2.3 All syringes, needles, lancets or other bloodletting devices capable of transmitting infection must be used only once and discarded in puncture resistant containers that are not overfilled. Sharps containers should be clearly marked to warn handlers of the potential hazard and should be located in areas where sharps are commonly used.
12.13 | Fire safety Is fire safety included as part of the laboratory’s overall safety programme? |  | 2 | Answer all sub questions | 0 |  | ISO 15190: 19.7 and 9.3 Electrical cords and plugs, power-strips and receptacles should be maintained in good condition and utilized appropriately. Overcrowding should be avoided and cords should be kept out of walkway areas. An approved fire extinguisher should be easily accessible within the laboratory and be routinely inspected and documented for readiness. Fire extinguishers should be kept in their assigned place and not hidden or blocked; the pin and seal should be intact, nozzles should be free of blockage, pressure gauges should show adequate pressure, and there should be no visible signs of damage. A fire alarm should be installed in the laboratory and tested regularly for readiness; all staff should participate in periodic fire drills.
|  | a) Are all electrical cords, plugs and receptacles used appropriately and in good repair? | [5 empty] | 100
|  | b) Is an appropriate fire extinguisher available, properly placed, in working condition and routinely inspected? | [5 empty] | 100
|  | c) Is an operational fire warning system in place? | [5 empty] | 100
|  | d) Are periodic fire drills conducted at defined intervals? | [5 empty] | 100 | 400
12.14 | Safety Audits Are safety inspections or audits conducted regularly and documented? |  | 3 | Answer all sub questions | 0 |  | ISO15190 Clause 7.3.1 and 7.3.2 Note: The safety programme shall be audited and reviewed at least annually (by appropriately trained personnel.
|  | a) Is there an audit plan/schedule that ensures all activities of the lab are checked for safety compliance? | [5 empty] | 100
|  | b) Are inspections/audits being carried out by authorized persons? | [5 empty] | 100
|  | c) Are the personnel conducting the internal audits trained in safety? | [5 empty] | 100
|  | d) Is root cause analysis performed and action taken for nonconformities/noted deficiencies? | [5 empty] | 100
|  | e) Are safety findings documented and presented to the laboratory management and relevant staff for review? | [5 empty] | 100 | 500
12.15 | Safety Equipment Is standard safety equipment available and in use in the laboratory? |  | 2 | Answer all sub questions | 0 |  | ISO15190 Clause 5.1 Note: It is the responsibility of laboratory management to ensure that the laboratory is equipped with standard safety equipment. The list above is a partial list of necessary items. Biosafety cabinets should be in place and in use as required. All centrifuges should have covers. Hand-washing stations should be designated and equipped and eyewash stations (or an acceptable alternative method of eye cleansing) should be available and operable. Spill kits and first aid kits should be kept in a designated place and checked regularly for readiness.
|  | a) Biosafety cabinet(s) |  |  |  |  | ISO 15190:2003 Clause 16 | 100
|  | b) Covers, safety caps, safety buckets on centrifuge(s) | [5 empty] | 100
|  | c) Hand-washing station |  |  |  |  | ISO 15190:2003 Clause 12.7 | 100
|  | d) Eyewash station/bottle(s) and emergency showers where applicable |  |  |  |  | ISO 15190:2003 Clause 12.10 | 100
|  | e) Spill kit(s) | [5 empty] | 100
|  | f) First aid kit(s) |  |  |  |  | ISO 15190:2003 Clause 12.9 | 100 | 600
12.16 | Personal Protective Equipment |  | 2 | Answer all sub questions | 0 |  | ISO 15190: 12 Management is responsible for providing appropriate personal protective equipment (gloves, lab coats, eye protection, etc) in useable condition. Laboratory staff must utilize PPE at all times while in the laboratory. Protective clothing should not be worn outside the laboratory. Gloves should be replaced immediately when torn or contaminated and not washed for reuse.
|  | Is personal protective equipment (PPE) easily accessible at the workstation and utilized appropriately and consistently? | [5 empty] | 100
|  | Answer to AMR question G12.1 | [5 empty] | 100 | 200
12.17 | Staff Vaccinations Are laboratory personnel offered appropriate vaccination and employee medical surveillance? |  | 2 |  | 0 |  | ISO 15190: 11.3 Laboratory staff should be offered appropriate vaccinations—particularly Hepatitis B. Staff may decline to receive the vaccination, but they must then sign a declination form to be held in the staff member’s personnel file.
12.18 | Post Exposure Prophylaxis |  | 2 | Answer all sub questions | 0 |  | ISO 15190: 9 The laboratory must have a procedure for follow-up of possible and known percutaneous, mucus membrane or abraded skin exposure to HIV, HBV or HCV. The procedure should include clinical and serological evaluation and appropriate prophylaxis.
|  | Are post-exposure prophylaxis policies and procedures posted and implemented after possible and known exposures? | [5 empty] | 100
|  | Answer to AMR question B12.2 | [5 empty] | 100 | 200
12.19 | Are adverse incidents or injuries from equipment, reagents, occupational injuries, medical screening or illnesses, documented and investigated? |  | 2 |  | 0 |  | ISO15189:2012 Clause 5.3.1.6; 5.3.2.6; ISO15190 Clause 9 Note: All occupational injuries or illnesses should be thoroughly investigated and documented in the safety log or occurrence log, depending on the laboratory. Corrective actions taken by the laboratory in response to an accident or injury must also be documented.
12.20 | Biosafety Training Are drivers/couriers and cleaners working with the laboratory trained in biosafety practices relevant to their job tasks? |  | 2 |  | 0 |  | ISO 15190: 10 All occupational injuries or illnesses should be thoroughly investigated and documented in the safety log or occurrence log, depending on the laboratory. Corrective actions taken by the laboratory in response to an accident or injury must also be documented.
12.21 | Laboratory Safety Officer Is a trained safety officer designated to implement and monitor the safety programme in the laboratory, including the training of other staff? |  | 2 |  | 0 |  | ISO 15190: 7, 10 A safety officer should be designated to work with the laboratory manager to implement the safety programme, monitor the ongoing safety conditions and needs of the laboratory, coordinate safety training, and serve as a resource for other staff. This officer should receive safety training.
| Totals |  | 43 |  | 0
Yes
Partial
No
N/A